## Instructions
July 2023
Instructions for using the QIAseq® UPXome RNA Library Prep with UDIs Template for the Local Run Manager® (LRM®) v2 Excel® File from Illumina®
Version 1.0 ¦ Adapted from the QIAseq UPXome RNA Library Kit Handbook
Tab Descriptions:
Tab 1 (Experiment Plan) This tab serves as space to document project and experimental details. It is the only sheet where data can be entered or changed outside of highlighted yellow cells.
Tab 2 (Sample Information Tab) Fill in the Sample ID, RNA Conc. and the desired ng per sample (E38) input. "0.1X QIAseq FastSelect dilution" will then be automatically calculated and populated into "FastSelect rRNA depletion procedure" found in the UPXome Library Construction tab. From the drop-down menus under RT Plate Location/SID Well ID (starting at G40), select which SID ID will be used for each sample; as a note, do not duplicate SID IDs as this will prohibit individual samples from being demultiplexed (if this occurs, a red font will be displayed, indicating SID IDs have been duplicated). From the drop-down menus under Pool Group Name (starting at H40), select which pool the samples will be in.
Tab 3 (UPXome Library Construction) This tab provides library construction guidance. The Pool ID will automatically autopopulate from the Sample Information tab. From the drop-down menus under QIAseq UX Index Plate Well Coordinate (in i7 RUDI, starting at E20), select which QIAseq UX library index will be used for each of the pools during "Library amplification/indexing procedure". The well ID for i5 RUDI will autopopulate. Do not duplicate QIAseq UX library indexes, as this will prohibit pools from being demultiplexed (if this occurs, a red font will be displayed, indicating QIAseq UX library indexes have been duplicated). As a note, master mix calculations include at least 10% extra.
Tab 4 (FastSelect Bacterial - Appx) This tab provides an alternative protocol when using FastSelect Bacterial. Once rRNA removal has been completed, continue with "cDNA synthesis procedure" in the UPXome Library Construction tab. As a note, master mix calculations include at least 10% extra.
Tab 5 (Dilute Lib) In this tab, copy and paste the concentrations in ng/µl from the Qubit results to the yellow section of the Dilute Lib tab, starting at Cell C43. After this, dilution for the samples to 4 nM will be automatically calculated. Add 3 µl of the library to the appropriate amount of water or HT1 Buffer. If your samples are too dilute, you may change cell D38 to be less than 4 nM, and the calculations will be automatically calculated for that concentration.
Tabs 6-8 (Denat Dilut Lib MiSeq, Denat Dilut Lib MiniSeq, or Denat Dilut Lib NextSeq) These tabs describe how to dilute the NaOH appropriately, how to denature and dilute pooled samples, and then how to load either a MiSeq®, MiniSeq®, or NextSeq® cartridge with 10% PhiX.
Tab 9 (Import Sample Sheet) This tab is autopopulated with information from the Sample Information tab. Fill the highlighted yellow cells, then copy this tab into a new book, remove and delete any excess rows and save it as a CSV file, which can then be saved where the sample sheets are saved on the Illumina instrument. The default sequencing structure is paired-end 2 x 74 bp with dual 10 bp sample indexes.
Note: Saving CSV files on operating systems that use commas as decimal separators may require you to replace all semicolons with commas using Notepad or any other text editor. Only sample sheets with commas as the filed separator can be read by Illumina instruments.
Tab 10 (CLC DEMUX and SID-ID) This tab contains the information needed by CLC to properly demultiplex the SID-IDs. Copy this tab into a new book, delete rows for all unused wells and any other excess rows, and then save it as a CSV file, which can then be used and imported into the CLC Quantify QIAseq UPXome workflow for demultiplexing individual samples.
Note: Saving CSV files on operating systems that use commas as decimal separators may require you to replace all semicolons with commas using Notepad or any other text editor. Only sample sheets with commas as the filed separator can be read by CLC.
Tab 11 (SID Plate Layout) This tab describes and illustrates the positions of the barcodes within the SID plates.
Tab 12 (QIAseq UX Index Plates) This tab describes and illustrates the positions of each sample index within the QIAseq UX index plates.
Tab 13 (UPXome library structure) This tab depicts the various QIAseq UPXome library structures.
QIAGEN shall not be held liable in connection with the use of this tool by any user. Proper testing/validation by the respective user is recommended prior to its use. By using this tool, you expressly accept the aforementioned exclusion of QIAGEN's liability.
For up-to-date licensing information and product-specific disclaimers, see the respective QIAGEN kit handbook or user manual. QIAGEN kit handbooks and user manuals are available at www.qiagen.com or can be requested from QIAGEN Technical Services or your local distributor.
Trademarks: QIAGEN®, Sample to Insight®, QIAseq®, QIAxcel® (QIAGEN Group); Bioanalyzer®, TapeStation® (Agilent Technologies, Inc.); Illumina®, Local Run Manager®, LRM®, MiSeq®, MiniSeq®, NextSeq® (Illumina, Inc.); Excel® (Microsoft Corporation). Registered names, trademarks, etc. used in this document, even when not specifically marked as such, are not to be considered unprotected by law.
03/2022 PROM-XXXXXX-001 © 2020 QIAGEN, all rights reserved.
Ordering www.qiagen.com/shop ¦ Technical Support support.qiagen.com ¦ Website www.qiagen.com

## 1. Experiment Plan
Experiment Plan
Experiment: | Initials: | Date:

## 2. Sample Information
QIAseq UPXome RNA Library Prep with UDIs for Illumina Instruments with Local Run Manager v2
Experiment: |  |  |  | Initials: |  |  | Date:
QIAseq UPXome RNA Library Prep with UDIs Sample Information Table
[_] | 1. Fill in Sample ID, RNA conc., and the ng per sample input amount in cell E38 for the samples being tested.
[_] | 2. The dilutions will then be automatically calculated for the indicated input amount to go directly into the FastSelect RNA removal
and Fragmentation reaction.
[_] | 3. From the drop-down menus under RT Plate Location SID Well ID (starting at G40), select which well ID will be used for each samples.
Important: Do not duplicate SID well IDs. Cells will display red font if SID well IDs are duplicated as a warning.
[_] | 4. From the drop-down menus under Pool Group Name (starting at H40), select which pool the samples will be in.
Comment: The volume amount (µl) of RNA to be added per sample and water to be added, as well as the
well ID to use, is autopopulated in the Summary Sample Table below.
Sample Summary Table
1. The table below shows how to dilute the samples appropriately for the "FastSelect rRNA depletion procedure" in the UPXome Library Construction tab.
ng per sample
[6 empty] | RT Plate Location | Pool Group Name
| Sample | Sample Name | RNA conc. (ng/µl) | RNA vol. (µl) | Water (µl) | SID Well ID
| 1 |  |  | #DIV/0! | #DIV/0!
| 2 |  |  | #DIV/0! | #DIV/0!
| 3 |  |  | #DIV/0! | #DIV/0!
| 4 |  |  | #DIV/0! | #DIV/0!
| 5 |  |  | #DIV/0! | #DIV/0!
| 6 |  |  | #DIV/0! | #DIV/0!
| 7 |  |  | #DIV/0! | #DIV/0!
| 8 |  |  | #DIV/0! | #DIV/0!
| 9 |  |  | #DIV/0! | #DIV/0!
| 10 |  |  | #DIV/0! | #DIV/0!
| 11 |  |  | #DIV/0! | #DIV/0!
| 12 |  |  | #DIV/0! | #DIV/0!
| 13 |  |  | #DIV/0! | #DIV/0!
| 14 |  |  | #DIV/0! | #DIV/0!
| 15 |  |  | #DIV/0! | #DIV/0!
| 16 |  |  | #DIV/0! | #DIV/0!
| 17 |  |  | #DIV/0! | #DIV/0!
| 18 |  |  | #DIV/0! | #DIV/0!
| 19 |  |  | #DIV/0! | #DIV/0!
| 20 |  |  | #DIV/0! | #DIV/0!
| 21 |  |  | #DIV/0! | #DIV/0!
| 22 |  |  | #DIV/0! | #DIV/0!
| 23 |  |  | #DIV/0! | #DIV/0!
| 24 |  |  | #DIV/0! | #DIV/0!
| 25 |  |  | #DIV/0! | #DIV/0!
| 26 |  |  | #DIV/0! | #DIV/0!
| 27 |  |  | #DIV/0! | #DIV/0!
| 28 |  |  | #DIV/0! | #DIV/0!
| 29 |  |  | #DIV/0! | #DIV/0!
| 30 |  |  | #DIV/0! | #DIV/0!
| 31 |  |  | #DIV/0! | #DIV/0!
| 32 |  |  | #DIV/0! | #DIV/0!
| 33 |  |  | #DIV/0! | #DIV/0!
| 34 |  |  | #DIV/0! | #DIV/0!
| 35 |  |  | #DIV/0! | #DIV/0!
| 36 |  |  | #DIV/0! | #DIV/0!
| 37 |  |  | #DIV/0! | #DIV/0!
| 38 |  |  | #DIV/0! | #DIV/0!
| 39 |  |  | #DIV/0! | #DIV/0!
| 40 |  |  | #DIV/0! | #DIV/0!
| 41 |  |  | #DIV/0! | #DIV/0!
| 42 |  |  | #DIV/0! | #DIV/0!
| 43 |  |  | #DIV/0! | #DIV/0!
| 44 |  |  | #DIV/0! | #DIV/0!
| 45 |  |  | #DIV/0! | #DIV/0!
| 46 |  |  | #DIV/0! | #DIV/0!
| 47 |  |  | #DIV/0! | #DIV/0!
| 48 |  |  | #DIV/0! | #DIV/0!
| 49 |  |  | #DIV/0! | #DIV/0!
| 50 |  |  | #DIV/0! | #DIV/0!
| 51 |  |  | #DIV/0! | #DIV/0!
| 52 |  |  | #DIV/0! | #DIV/0!
| 53 |  |  | #DIV/0! | #DIV/0!
| 54 |  |  | #DIV/0! | #DIV/0!
| 55 |  |  | #DIV/0! | #DIV/0!
| 56 |  |  | #DIV/0! | #DIV/0!
| 57 |  |  | #DIV/0! | #DIV/0!
| 58 |  |  | #DIV/0! | #DIV/0!
| 59 |  |  | #DIV/0! | #DIV/0!
| 60 |  |  | #DIV/0! | #DIV/0!
| 61 |  |  | #DIV/0! | #DIV/0!
| 62 |  |  | #DIV/0! | #DIV/0!
| 63 |  |  | #DIV/0! | #DIV/0!
| 64 |  |  | #DIV/0! | #DIV/0!
| 65 |  |  | #DIV/0! | #DIV/0!
| 66 |  |  | #DIV/0! | #DIV/0!
| 67 |  |  | #DIV/0! | #DIV/0!
| 68 |  |  | #DIV/0! | #DIV/0!
| 69 |  |  | #DIV/0! | #DIV/0!
| 70 |  |  | #DIV/0! | #DIV/0!
| 71 |  |  | #DIV/0! | #DIV/0!
| 72 |  |  | #DIV/0! | #DIV/0!
| 73 |  |  | #DIV/0! | #DIV/0!
| 74 |  |  | #DIV/0! | #DIV/0!
| 75 |  |  | #DIV/0! | #DIV/0!
| 76 |  |  | #DIV/0! | #DIV/0!
| 77 |  |  | #DIV/0! | #DIV/0!
| 78 |  |  | #DIV/0! | #DIV/0!
| 79 |  |  | #DIV/0! | #DIV/0!
| 80 |  |  | #DIV/0! | #DIV/0!
| 81 |  |  | #DIV/0! | #DIV/0!
| 82 |  |  | #DIV/0! | #DIV/0!
| 83 |  |  | #DIV/0! | #DIV/0!
| 84 |  |  | #DIV/0! | #DIV/0!
| 85 |  |  | #DIV/0! | #DIV/0!
| 86 |  |  | #DIV/0! | #DIV/0!
| 87 |  |  | #DIV/0! | #DIV/0!
| 88 |  |  | #DIV/0! | #DIV/0!
| 89 |  |  | #DIV/0! | #DIV/0!
| 90 |  |  | #DIV/0! | #DIV/0!
| 91 |  |  | #DIV/0! | #DIV/0!
| 92 |  |  | #DIV/0! | #DIV/0!
| 93 |  |  | #DIV/0! | #DIV/0!
| 94 |  |  | #DIV/0! | #DIV/0!
| 95 |  |  | #DIV/0! | #DIV/0!
| 96 |  |  | #DIV/0! | #DIV/0!
| Number of samples: |  |  |  | 0

## 3. UPXome Library Construction
QIAseq UPXome RNA Library Prep with UDIs for Illumina Instruments with Local Run Manager v2
Experiment: |  |  |  | Initials: |  |  | Date:
QIAseq UPXome RNA Library Prep with UDIs Sample Information Table
Sample Summary Table
[_] | 1. The table below shows the Pool ID that is autopopulated from the Sample Information tab.
[_] | 2. From the drop-down menus under QIAseq UX Index Plate Well Coordinate (starting at E20), select the coordinate of the i7 RUDI index for each of the pooled samples in Library Amplification. Do not duplicate QIAseq UX library indexes, as this will prohibit pools from being demultiplexed.
QIAseq UX Index Plate Well Coordinate
|  |  | Pool ID | i7 RUDI | i5 RUDI
|  |  | 0 |  | 0
0
0
0
0
0
0
0
0
0
0
0
| Number of samples: |  |  |  | 0
The procedures below are the default protocol for QIAseq UPXome Library Construction. If using QIAseq FastSelect Bacterial –5S/16S/23S, first proceed to the tab "FastSelect Bacterial - Appx".
FastSelect rRNA depletion procedure
Important points before starting:
[_] | 1. This protocol can be used with low amounts (500 pg – 100 ng) of purified RNA. The recommended starting amount is 10 ng (or greater).
[_] | 2. After cDNA synthesis, at least eight samples and no more than 24 samples must be pooled together.
[_] | 3. DNase treatment (on-column or in-solution) of total RNA samples is highly recommended.
[_] | 4. ERCC Control RNA can be added according to the concentrations specified by the manufacturer. If added, the total reaction volume should remain 11 µl.
[_] | 5. Gently yet thoroughly vortex reactions or reagents unless instructed otherwise. Ensure all reactions are mixed thoroughly and handled at the temperatures recommended in the protocol.
[_] | 6. Use a thermal cycler with a heated lid.
[_] | 7. Thaw template RNA on ice.
[_] | 8. Thaw and dilute the QIAseq FastSelect to 0.1X as follows:
|  |  | QIAseq FastSelect | 0 | [_]
|  |  | Water | 0 | [_]
[_] | 9. Thaw the US RT Buffer, 5x and Nuclease-free Water at room temperature. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
Procedure
[_] | 1. Prepare the FastSelect rRNA depletion mix (0.1X QIAaseq FastSelect + US RT Buffer, 5x) as follows:
Volume (µl)
|  |  | Reagents |  | 1x | 0
|  |  | RNA (value from Sample Summary Table)* |  | Variable | - | [_]
|  |  | Nuclease-free Water (value from Sample Summary Table)* |  | Variable | - | [_]
|  |  | 0.1X QIAseq FastSelect |  | 1 | 0 | [_]
|  |  | US RT Buffer, 5x |  | 4 | 0 | [_]
|  |  | Total volume |  | 5 | 0
* These values have been automatically calculated in the Sample Summary Table (Sample Information tab).
[_] | 2. Add 5 µl of the FastSelect rRNA depletion mix to the 6 µl "RNA + Nuclease-free Water" from the Sample Summary Table (11 µl total).
[_] | 3. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
[_] | 4. Incubate as follows using a thermal cycler with a heated lid:
|  | Step | Time and Temperature
|  | 1 | 2 min at 75°C
|  | 2 | 2 min at 70°C
|  | 3 | 2 min at 65°C
|  | 4 | 2 min at 60°C
|  | 5 | 2 min at 55°C
|  | 6 | 2 min at 37°C
|  | 7 | 2 min at 25°C
|  | 8 | 2 min at 4°C
|  | 9 | Hold at 4°C
[_] | 5. Proceed to "cDNA synthesis procedure".
cDNA synthesis procedure
Important points before starting:
[_] | 1. Set up the reactions on ice.
[_] | 2. Use a thermal cycler with a heated lid.
[_] | 3. The SID-TS-96S or SID-TS-24S is a 96-well single-use sample ID RT plate. Each well of the SID-TS-96S and only 24 wells of the SID-TS-24S Plate contain pre-dispensed sample ID template switching oligos required for the cDNA synthesis reaction.
Important: If, during the setup of the cDNA synthesis reactions, only some of the wells in the 96‑well plate will be used, add 2.5 µl Nuclease-free Water to each well intended to be used, vortex the plate, centrifuge briefly, and incubate for 10 min at room temperature (15–25°C) to fully dissolve the primer. Then transfers the dissolved content of wells that are going to be used to a new plate. Unused wells should be sealed with provided strip caps.
[_] | 4. Equilibrate the QIAseq Beads to room temperature for 20–30 min before use.
[_] | 5. Vortex the QIAseq Beads thoroughly to ensure they are well resuspended; the solution should be a homogenous brown color. If a delay in the protocol occurs, simply vortex the beads.
[_] | 6. Important: Prepare fresh 80% ethanol before performing the procedure.
[_] | 7. Thaw DTT (100 mM), dNTP (10 mM), Nuclease-free Water, N6-T RT Primer and/or ODT-T RT Primer at room temperature (15-25ºC). Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
[_] | 8. Keep RNase Inhibitor and EZ Reverse Transcriptase on ice. Mix by flicking the tubes, briefly centrifuge and return to ice. After use, immediately return the enzymes to the freezer.
Procedure
[_] | 1. Prepare the cDNA synthesis reaction mix on ice as follows:
Volume (µl)
| Reagents |  |  | 1x | 0
| rRNA-depleted sample from previous step |  |  | 11 | - | [_]
| DTT (100 mM) |  |  | 0.5 | 0 | [_]
| dNTP (10 mM) |  |  | 2 | 0 | [_]
| Nuclease-free Water* |  |  | 2.5 | 0 | [_]
| SID-TS RT Plate Dried* |  |  | - | - | [_]
| N6-T RT Primer† |  |  | 1 | 0 | [_]
| ODT-T RT Primer† |  |  | 1 | 0 | [_]
| EZ Reverse Transcriptase |  |  | 1.5 | 0 | [_]
| RNase Inhibitor |  |  | 0.5 | 0 | [_]
| Total volume |  |  | 20 | 0
* If not using all the wells in the SID-TS-96S or SID-TS-24S Plate, reduce the 2.5 µl volume of Nuclease-free Water to 0 µl in the cDNA synthesis mix, as 2.5 µl of Nuclease-free water will be used to resuspend the sample ID template switching oligos that will be used.
† For 3’ transcriptome analysis, only use the ODT-T RT primer, and replace the N6-T RT primer with an additional 1 µl of Nuclease-free Water.
[_] | 2. Add the cDNA synthesis mix to the corresponding wells in the SID-TS-96S plate or SID-TS-24S.
Important: A unique sample ID must be used for each sample.
[_] | 3. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
[_] | 4. Incubate as follows using a thermal cycler with a heated lid:
|  | Step | Time and Temperature
|  | 1 | 1 min at 4°C
|  | 2 | 5 min at 25°C
|  | 3 | 90 min at 42°C
|  | 4 | 10 min at 70°C
|  | 5 | 1 min at 4°C
|  | 6 | Hold at 4°C
[_] | 5. Upon completion of the cDNA synthesis reactions, pool at least eight samples and no more than 24 samples together into one 1.7 ml (or 2.0 ml) microcentrifuge tube.
Note: The cDNA generated from each well of a SID-TS RT Plate is barcoded.
[_] | 6. Add the appropriate volume (1.1x the reaction volume) of QIAseq Beads to the combined cDNA synthesis reactions from Step 5 according to the following table. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
Note: Vortex the QIAseq Beads thoroughly to ensure they are well resuspended; the solution should be a homogenous brown color.
Volume (µl)
|  | Number of wells combined |  | Sample | QIAseq Bead (Ratio 1.1x)
|  | 8 |  | 160 | 176
|  | 24 |  | 480 | 528
[_] | 7. Incubate for 5 min at room temperature.
[_] | 8. Centrifuge the tube until the beads are pelleted (2 min), and then place the plate onto a magnetic rack. After the solution has cleared (2 min or longer), carefully remove and discard the supernatant.
Important: Do not discard the beads, as they contain the DNA of interest.
Tip: It may be valuable to discard the supernatant twice. The contents settle after the first discard.
[_] | 9. With the tube still on the magnetic stand, add 200−300 µl freshly prepared 80% ethanol. Wait 2 min and carefully remove and discard the wash. Repeat the wash, for a total of 2 ethanol washes. Remove as much excess ethanol as possible.
Note: When decided between 200 µl or 300 µl, simply ensure that enough is added to submerge the beads.
Important: It is vital to completely remove all traces of the ethanol after the second wash. Remove the ethanol with a 200 µl pipette first, briefly centrifuge the tubes immediately and return the tubes to the magnetic stand. Then use a 10 µl pipette to remove any residual ethanol. This step should be performed quickly.
[_] | 10. With the tube (cap open) still on the magnetic stand, air-dry at room temperature for 5–10 min (until the beads start to crack and pellet loses its shine).
Note: Visually inspect that the pellet is completely dry.
[_] | 11. Remove the beads from the magnetic stand and elute the DNA from the beads by adding 22 µl Nuclease-free Water. Gently yet thoroughly vortex (triturate if necessary) to mix, briefly centrifuge and incubate for 2 min.
[_] | 12. Centrifuge the tube until the beads are pelleted (2 min), and then return the tube to the magnetic rack until the solution has cleared.
[_] | 13. Transfer 20 µl of the supernatant to a clean tube.
[_] | 14. Add 22 µl of QIAseq Beads (1.1x the reaction volume) to the supernatant. Briefly centrifuge, gently yet thoroughly vortex to mix and centrifuge briefly again.
Note: Vortex the QIAseq Beads thoroughly to ensure they are well resuspended; the solution should be a homogenous brown color.
[_] | 15. Incubate for 5 min at room temperature.
[_] | 16. Centrifuge the tube until the beads are pelleted (2 min), and then place the plate onto a magnetic rack. After the solution has cleared (2 min or longer), carefully remove and discard the supernatant.
Important: Do not discard the beads, as they contain the DNA of interest.
Tip: It may be valuable to discard the supernatant twice. The contents settle after the first discard.
[_] | 17. With the tube still on the magnetic stand, add 200 µl freshly prepared 80% ethanol. Wait 2 min and carefully remove and discard the wash. Repeat the wash, for a total of two ethanol washes. Remove excess ethanol as much as possible.
Important: It is vital to completely remove all traces of the ethanol after the second wash. Remove the ethanol with a 200 µl pipette first, briefly centrifuge the tubes immediately and return the tubes to the magnetic stand. Then use a 10 µl pipette to remove any residual ethanol. This step should be performed quickly.
[_] | 18. With the tube (cap open) still on the magnetic stand, air-dry at room temperature for 5–10 min (until the beads start to crack and pellet loses its shine).
Note: Visually inspect that the pellet is completely dry.
[_] | 19. Remove the beads from the magnetic stand and elute the DNA from the beads by adding 25 µl Nuclease-free Water. Gently yet thoroughly vortex (triturate if necessary) to mix, briefly centrifuge, and incubate for 2 min.
[_] | 20. Centrifuge the tube until the beads are pelleted (2 min), and then return the tube to the magnetic rack until the solution has cleared.
[_] | 21. Transfer 23 µl of the supernatant to a clean tube.
Note: From this point forward in the protocol, the procedures are written with the assumption that all cDNA wells (either 8 or 24) have been combined into a single tube. In your experiment, you may be using more tubes.
[_] | 22. Proceed with “Library amplification/indexing procedure". Alternatively, the samples can be stored at −30°C to −15°C.
Library amplification/indexing procedure
Important points before starting:
[_] | 1. The QIAseq UX index plates have pierceable foil seals, and the indexes must be pipetted from the plate into separate reaction plates. To prevent cross-contamination, each well is single use.
[_] | 2. Set up the reactions on ice.
[_] | 3. Use a thermal cycler with a heated lid.
[_] | 4. Equilibrate the QIAseq Beads to room temperature (15–25°C) for 20–30 min before use.
[_] | 5. Vortex the QIAseq Beads thoroughly to ensure they are well resuspended; the solution should be a homogenous brown color. If a delay in the protocol occurs, simply vortex the beads.
[_] | 6. Important: Prepare fresh 80% ethanol before performing this procedure.
[_] | 7. Thaw the QIAseq UX index plates at room temperature (15–25°C). Gently yet thoroughly vortex to mix, and then centrifuge briefly.
[_] | 8. Thaw the QIAseq 2X HiFi MM on ice. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
Note: A precipitate might be present in the QIAseq 2X HiFi MM. If so, bring the reagent to room temperature for 5 min, and dissolve the precipitate by mixing with a pipettor and/or gentle vortexing.
Procedure
[_] | 1. For each set of pooled cDNA, prepare the library amplification/indexing reaction mix on ice as follows:
Volume (µl)
|  | Reagents |  | 1x
|  | Pooled cDNA |  | 23 | [_]
|  | QIAseq 2X HiFi MM |  | 25 | [_]
|  | Unique index from QIAseq UX index plate* |  | 2 | [_]
|  | Total volume |  | 50
* Important: For each pooled cDNA sample, ensure a unique QIAseq UX index has been used.
[_] | 2. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
[_] | 3. In a thermal cycler, perform the following cycling program. The number of cycles is determined by the subsequent table:
| Step |  | Temperature and time |  | Number of cycles
| Initial denaturation |  | 98°C for 30 s |  | 1
| PCR |  | 98°C for 5 s |  | Refer to subsequent table
55°C for 10 s
72°C for 20 s
| Final extension |  | 72°C for 2 min |  | 1
| Hold |  | 4°C for 1 min |  | 1
|  |  | 4°C for ∞ |  | HOLD
Number of amplification cycles
| Total RNA input |  | cDNA pool of eight samples |  | cDNA pool of 24 samples
| 500 pg |  | 24 |  | 23
| 1 ng |  | 20 |  | 19
| 10 ng |  | 18 |  | 17
| 100 ng |  | 15 |  | 14
[_] | 4. Upon completion of library-amplification or indexing reactions, add 40 µl QIAseq Beads (0.8x the reaction volume). Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
Note: Vortex the QIAseq Beads thoroughly to ensure they are well resuspended; the solution should be a homogenous brown color.
[_] | 5. Incubate for 5 min at room temperature.
[_] | 6. Centrifuge the tube until the beads are pelleted (2 min), and then place the plate onto a magnetic rack. After the solution has cleared (2 min or longer), carefully remove and discard the supernatant.
Important: Do not discard the beads, as they contain the DNA of interest
Tip: It may be valuable to discard the supernatant twice. The contents settle after the first discard
[_] | 7. With the tube still on the magnetic stand, add 200 µl freshly prepared 80% ethanol. Wait 2 min and carefully remove and discard the wash. Repeat the wash, for a total of 2 ethanol washes. Remove as much excess ethanol as possible.
Important: It is vital to completely remove all traces of the ethanol after the second wash. Remove the ethanol with a 200 µl pipette first, briefly centrifuge the tubes immediately and return the tubes to the magnetic stand. Then use a 10 µl pipette to remove any residual ethanol. This step should be performed quickly.
[_] | 8. With the tube (cap open) still on the magnetic stand, air-dry at room temperature for 5–10 min (until the beads start to crack and pellet loses its shine).
Note: Visually inspect that the pellet is completely dry.
[_] | 9. Remove the beads from the magnetic stand and elute the DNA from the beads by adding 22 µl Nuclease-free Water. Gently yet thoroughly vortex (triturate if necessary) to mix, briefly centrifuge and incubate for 2 min.
[_] | 10. Centrifuge the tube until the beads are pelleted (2 min), and then return the tube to the magnetic rack until the solution has cleared.
[_] | 11. Transfer 20 µl of the supernatant to a clean tube.
[_] | 12. Add 16 µl of QIAseq Beads (0.8x the reaction volume) to the supernatant. Briefly centrifuge, gently yet thoroughly vortex to mix and centrifuge briefly again.
Note: Vortex the QIAseq Beads thoroughly to ensure they are well resuspended; the solution should be a homogenous brown color.
[_] | 13. Incubate for 5 min at room temperature.
[_] | 14. Centrifuge the tube until the beads are pelleted (2 min), and then place the plate onto a magnetic rack. After the solution has cleared (2 min or longer), carefully remove and discard the supernatant.
Important: Do not discard the beads, as they contain the DNA of interest.
Tip: It may be valuable to discard the supernatant twice. The contents settle after the first discard.
[_] | 15. With the tube still on the magnetic stand, add 200 µl freshly prepared 80% ethanol. Wait 2 min and carefully remove and discard the wash. Repeat the wash, for a total of two ethanol washes. Remove excess ethanol as much as possible.
Important: It is vital to completely remove all traces of the ethanol after the second wash. Remove the ethanol with a 200 µl pipette first, briefly centrifuge the tubes immediately and return the tubes to the magnetic stand. Then use a 10 µl pipette to remove any residual ethanol. This step should be performed quickly.
[_] | 16. With the tube (cap open) still on the magnetic stand, air-dry at room temperature for 5–10 min (until the beads start to crack and pellet loses its shine).
Note: Visually inspect that the pellet is completely dry.
[_] | 17. Remove the beads from the magnetic stand and elute the DNA from the beads by adding 24 µl Nuclease-free Water. Gently yet thoroughly vortex (triturate if necessary) to mix, briefly centrifuge, and incubate for 2 min.
[_] | 18. Centrifuge the tube until the beads are pelleted (2 min), and then return the tube to the magnetic rack until the solution has cleared.
[_] | 19. Transfer 22 µl of the supernatant to a clean tube. If not proceeding immediate, the samples can be stored at −30°C to −15°C.
[_] | 20. Proceed to assess the quality and quantity of the libraries. Qualify the libraries using either a QIAxcel® (QIAGEN), Bioanalyzer® (Agilent), or TapeStation® (Agilent) measurements. Check for the expected size distribution of approximately 300–500 bp, adapter-dimer-free (nothing at 130 bp). Quantify the libraries using a Qubit Fluorimeter (Thermo Fisher).

## 4. FastSelect Bacterial - Appx
Appendix: QIAseq FastSelect –5S/16S/23S with the QIAseq UPXome RNA Library Kit
If using QIAseq FastSelect Bacterial –5S/16S/23S, proceed as described below.
[_] | 1. This protocol is for QIAseq FastSelect –5S/16S/23S rRNA removal and any QIAseq FastSelect combination with QIAseq FastSelect –5S/16S/23S and the starting material is 10 ng to 100 ng total RNA. Using less than 10 ng input is not recommended when using QIAseq FastSelect –5S/16S/23S.
[_] | 2. After cDNA synthesis, at least eight samples and no more than 24 samples must be pooled together.
[_] | 3. DNase treatment (on-column or in-solution) of total RNA samples is highly recommended.
[_] | 4. ERCC Control RNA can be added according to the concentrations specified by the manufacturer. If added, the total reaction volume should remain 15 µl.
[_] | 5. Gently yet thoroughly vortex reactions or reagents unless instructed otherwise. Ensure all reactions are mixed thoroughly and handled at the temperatures recommended in the protocol.
[_] | 6. Equilibrate the QIAseq Beads and QIAseq Bead Binding Buffer to room temperature for 20–30 min before use.
[_] | 7. Use a thermal cycler with a heated lid.
[_] | 8. Thaw template RNA on ice.
[_] | 9. Thaw and dilute the QIAseq FastSelect 5S/16S/23S to 0.1X as follows:
Important: Only for QIAseq FastSelect 5S/16S/23S, incubate the tube at 37°C for 5 min, and then vortex to dissolve the precipitate.
|  | QIAseq FastSelect 5S/16S/23S |  | 0 | [_]
|  | Water |  | 0 | [_]
[_] | 10. Thaw FastSelect FH Buffer and Nuclease-free Water at room temperature. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
Procedure
[_] | 1. Prepare the FastSelect rRNA depletion mix as follows:
Volume (µl)
|  | Reagents |  | 1x | 0
|  | RNA (value from Sample Summary Table)* |  | Variable | - | [_]
|  | Nuclease-free Water (value from Sample Summary Table)* |  | Variable | - | [_]
|  | 0.1X QIAseq FastSelect 5S/16S/23S |  | 1 | 0 | [_]
|  | FastSelect FH Buffer |  | 1.5 | 0 | [_]
|  | Additional Nuclease-free Water |  | 6.5 | 0 | [_]
|  | Total volume |  | 15 | 0
* These values have been automatically calculated in the Sample Summary Table (Sample Information tab).
[_] | 2. Add 9 µl of the FastSelect rRNA depletion mix to the 6 µl "RNA + Nuclease-free Water" from the Sample Summary Table (15 µl total).
[_] | 3. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
[_] | 4. Incubate as follows using a thermal cycler with a heated lid:
|  | Step | Time and Temperature
|  | 1 | 2 min at 75°C
|  | 2 | 2 min at 70°C
|  | 3 | 2 min at 65°C
|  | 4 | 2 min at 60°C
|  | 5 | 2 min at 55°C
|  | 6 | 2 min at 37°C
|  | 7 | 2 min at 25°C
|  | 8 | 2 min at 4°C
|  | 9 | Hold at 4°C
[_] | 5. Upon completion of the rRNA depletion reaction, add 19.5 µl (1.3x volume) QIAseq Beads to the 15 µl reaction. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
Note: Vortex the QIAseq Beads thoroughly to ensure they are well resuspended; the solution should be a homogenous brown color.
[_] | 6. Incubate for 5 min at room temperature.
[_] | 7. Centrifuge the tube or plate until the beads are pelleted (2 min), and then place the plate onto a magnetic rack. After the solution has cleared (2 min or longer), carefully remove and discard the supernatant.
Important: Do not discard the beads, as they contain the DNA of interest.
Tip: It may be useful to discard the supernatant twice. The contents settle after the first discard.
[_] | 8. Add 15 µl of Nuclease-free Water and 19.5 µl QIAseq Bead Binding Buffer (prewarmed to room temperature and mixed by vortexing). Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
[_] | 9. Incubate for 5 min at room temperature.
[_] | 10. Centrifuge the tube or plate until the beads are pelleted (2 min), and then place the tube or plate onto a magnetic rack. After the solution has cleared (2 min or longer), carefully remove and discard the supernatant.
Important: Do not discard the beads, as they contain the DNA of interest.
Tip: It may be useful to discard the supernatant twice. The contents settle after the first discard.
[_] | 11. With the tube still on the magnetic stand, add 200 µl freshly prepared 80% ethanol. Wait 2 min and carefully remove and discard the wash. Repeat the wash, for a total of two ethanol washes. Remove as much excess ethanol as possible.
Important: It is vital to completely remove all traces of the ethanol after the second wash. Remove the ethanol with a 200 µl pipette first, briefly centrifuge the tubes immediately and return the tubes to the magnetic stand. Then use a 10 µl pipette to remove any residual ethanol. This step should be performed quickly.
[_] | 12. With the tube (cap open) still on the magnetic stand, air-dry at room temperature for 5–10 min (until the beads start to crack and pellet loses its shine).
Note: Visually inspect that the pellet is completely dry.
[_] | 13. Remove the beads from the magnetic stand and elute the DNA from the beads by adding 9 µl Nuclease-free Water. Gently yet thoroughly vortex (triturate if necessary) to mix, briefly centrifuge, and incubate for 2 min.
[_] | 14. Centrifuge the tube until the beads are pelleted (2 min), and then return the tube to the magnetic rack until the solution has cleared.
[_] | 15. Transfer 7 µl of the supernatant to a clean tube.
[_] | 16. Thaw the US RT Buffer, 5x and add 4 µl to the 7 µl supernatant. Briefly centrifuge, gently yet thoroughly vortex to mix, and centrifuge briefly again.
[_] | 17. Proceed to "cDNA synthesis procedure" in UPXome Library Construction tab. Alternatively, the samples can be stored at −30°C to −15°C in a constant-temperature freezer.

## 5. Dilute Lib
Diluting the QIAseq UPXome libraries for Illumina sequencing:
Important points before starting:
[_] | 1. A Qubit Fluorimeter is recommended to determine the library concentration.
Determining library concentration procedure
[_] | 1. Determine the concentration of 2 µl of the QIAseq UPXome Sequencing Library on a Qubit Fluorimeter according to the manufacturer’s instructions.
[_] | 2. Input the concentration in ng/µl in the table below, and it will automatically calculate the nM and then tell you how to dilute the samples to 4 nM.
[_] | 3. Dilute the individual libraries to 4 nM using nuclease-free water.
[_] | 4. If multiplexing, combine libraries in equimolar amounts and mix well.
|  |  | Average Library size: | 500
|  |  | Library (nM): | 4
|  | Sample | ng/µl | Sample [ ] (µM) | Sample [ ] (nM) | Dilution Factor | Lib Final conc (nM) | Volume Water
4
|  | 1 |  | 0 | 0 | 0 | 3 | -3
|  | 2 |  | 0 | 0 | 0 | 3 | -3
|  | 3 |  | 0 | 0 | 0 | 3 | -3
|  | 4 |  | 0 | 0 | 0 | 3 | -3
|  | 5 |  | 0 | 0 | 0 | 3 | -3
|  | 6 |  | 0 | 0 | 0 | 3 | -3
|  | 7 |  | 0 | 0 | 0 | 3 | -3
|  | 8 |  | 0 | 0 | 0 | 3 | -3
|  | 9 |  | 0 | 0 | 0 | 3 | -3
|  | 10 |  | 0 | 0 | 0 | 3 | -3
|  | 11 |  | 0 | 0 | 0 | 3 | -3
|  | 12 |  | 0 | 0 | 0 | 3 | -3
Diluting and preparing the pooled library procedure
[_] | 1. Dilute the sample libraries down to 4 nM using the calculations above. Add 3 μl of the library sample to the appropriate volume of water or HT1 buffer.
[_] | 2. Multiple samples are now combined to create the "Lane Pool" – combined lane at 4 nM. Therefore, take 5 μl of each sample and combine them in a tube. Mix thoroughly.

## SID Plate Locations (hidden)
ATGTCTTACG | A01 | QID-3 | Pool_ID_1
CAAGTAGTCT | B01 | QID-113 | Pool_ID_2
GGATAGGTCT | C01 | QID-44 | Pool_ID_3
TACACTACGA | D01 | QID-276 | Pool_ID_4
ATATTCGACG | E01 | QID-110 | Pool_ID_5
CAGTCCAATT | F01 | QID-4 | Pool_ID_6
GACAACTGTC | G01 | QID-24 | Pool_ID_7
TCAAGCAATC | H01 | QID-75 | Pool_ID_8
ATCGTGTTCT | A02 | QID-2 | Pool_ID_9
CATTCGCGTC | B02 | QID-5 | Pool_ID_10
GAGAGCTACT | C02 | QID-28 | Pool_ID_11
TGACAGCGTG | D02 | QID-339 | Pool_ID_12
ATGTAGGCCT | E02 | QID-111
CCGATGGTTG | F02 | QID-6
GCTTCCATTC | G02 | QID-202
TGGAACTCGG | H02 | QID-91
AAGTTCGTGC | A03 | QID-1
CGGTCTCTAT | B03 | QID-125
GAACAATCCA | C03 | QID-140
TTAACACGAG | D03 | QID-366
ATGTGGAGGT | E03 | QID-112
CTACCGGTTC | F03 | QID-128
GTACCATGCC | G03 | QID-54
TTAATGGAGG | H03 | QID-99
CTTCAATCCT | A04 | QID-17
GAGATAACTG | B04 | QID-29
TAATCTAGGC | C04 | QID-65
GCTTAGTCAT | D04 | QID-41
CTTGGTGATT | E04 | QID-18
TCCAGAGTCC | F04 | QID-77
GTAATAGCCA | G04 | QID-53
TGCCTAACGC | H04 | QID-89
TCGATTACCA | A05 | QID-78
GAGCAGCCTT | B05 | QID-30
CCGGTAAGCT | C05 | QID-7
TACCGTTCTT | D05 | QID-66
GGCATAGTGG | E05 | QID-45
CTTGTGTCGG | F05 | QID-19
TTGGCATTAC | G05 | QID-104
CTGGAGGATA | H05 | QID-134
CGGTTGTAGC | A06 | QID-127
GGACCTTATT | B06 | QID-43
TACGAGCTGC | C06 | QID-67
CATTGCATAG | D06 | QID-115
GTAGGAACTT | E06 | QID-55
TCGCATCACG | F06 | QID-79
CTTCTGACAT | G06 | QID-139
TTCTTCCAAC | H06 | QID-103
GTACACTCAT | A07 | QID-225
TGTCAGGCTC | B07 | QID-356
CATTGGTGCG | C07 | QID-116
GACCTTCGTC | D07 | QID-163
CCGTCTGAAC | E07 | QID-8
GCATATCCTT | F07 | QID-175
TGCCTGATTA | G07 | QID-343
GAGTAAGGAG | H07 | QID-31
TCGCGTGGTA | A08 | QID-80
CGGTTATCCG | B08 | QID-9
GAACTTGTTG | C08 | QID-20
CTACTCAGAC | D08 | QID-11
TCTGAGCCTT | E08 | QID-333
CCGCTACTTA | F08 | QID-119
GTTCATGGTA | G08 | QID-249
TACGGCAACG | H08 | QID-69
CGGTTCAGTG | A09 | QID-126
GAAGGCATCT | B09 | QID-152
TCGTCCGTCA | C09 | QID-81
CTCCTTGCCA | D09 | QID-131
TCACACAGTA | E09 | QID-309
GAAGTACAGT | F09 | QID-22
TGGAGTCTTG | G09 | QID-92
CTAGTGCTTA | H09 | QID-12
GCTATGAATC | A10 | QID-188
CTCGTCATCC | B10 | QID-132
TCTCGGTTCG | C10 | QID-82
CCGCCTTACT | D10 | QID-117
GCATGTTAGT | E10 | QID-176
CTGGTGTGCA | F10 | QID-135
TAGACACCAA | G10 | QID-70
GAAGCGTGAT | H10 | QID-21
GCAGTGATCG | A11 | QID-33
CTACCTCTCT | B11 | QID-10
TACGGTTAGA | C11 | QID-283
GTGTTACTGA | D11 | QID-57
TTGGTTAGTC | E11 | QID-106
CTATCCGCTG | F11 | QID-13
GCTCAGCTTG | G11 | QID-189
CATTGACTCT | H11 | QID-114
GAGGTGAACA | A12 | QID-165
CTATCCTAGC | B12 | QID-129
TGTCCTGTAC | C12 | QID-94
GAACACCACG | D12 | QID-141
CCGTCTTGTG | E12 | QID-121
TGCTAGTAGT | F12 | QID-90
GCTCATTGCT | G12 | QID-190
CTGTGGACAC | H12 | QID-15

## RUDI Plate A Layout (hidden)
QIAseq RUDI Adapter Plate A (96)

## Index Plates (hidden)
QIAseq RUDI Plate A (96) | QIAseq RUDI Plate B (96)
Cat. no. | Cat. no.

## 6. Denat Dilut Lib MiSeq
How to Denature and Dilute Libraries for the MiSeq
Experiment: |  |  | Initials: | Date:
MiSeq run information: |  |  |  | Date:
Denaturing and diluting the sample pool libraries to be loaded onto the
MiSeq instrument
Preparation of fresh 0.2 N NaOH
1. Prepare fresh 0.2 N NaOH from 10 N NaOH stock:
[_] | 1.a Add 20 μl of 10 N NaOH stock to 180 μl of nuclease-free water. Mix thoroughly.
The NaOH is now at a concentration of 1 N NaOH.
[_] | 1.b Add 100 μl of 1 N NaOH dilution from step 1.a above to 400 μl of nuclease-free water.
Mix thoroughly. The NaOH is now at a concentration of 0.2 N NaOH.
Denaturing and diluting the pooled library into position #17
[_] | 1. Take 5 μl of the combined pooled 4 nM lane from the Dilute Lib tab
and add it to 5 μl of 0.2 N NaOH. The samples are now at a concentration of 2 nM.
[_] | 2. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 3. Denature the 2 nM sample for 5 min at room temperature.
[_] | 4. Add 990 μl of the HT1 Hybridization Buffer to the 10 μl reaction from step 3. The samples are
now at a 20 pM concentration and are ready to be diluted prior to being loaded onto the MiSeq
Reagent Cartridge.
[_] | 5. Dilute the sample again to a final concentration appropriate for loading onto the MiSeq Reagent
Cartridge. We typically load a final concentration of 10 pM onto the MiSeq Reagent Cartridge
with 10% PhiX Control Spike-in into Sample Position #17. Hence, add 270 µl of HT1 Buffer,
30 µl of a denatured and diluted 20 pM PhiX Control to 300 µl of the 20 pM denatured
library for a final volume of 600 µl. Mix thoroughly before adding to Position #17.
[_] | 6. Discard the unused 0.2 N NaOH.
MiSeq run setup information:
| 1. Sequencing setup: | Select Category: Other > FASTQ Only
Read Type: Paired End Read
Cycles: 74 for R1, 74 for R2
Indices: 10 bp
2. Sequenced with 10% PhiX Control

## 7. Denat Dilut Lib MiniSeq
How to Denature and Dilute Libraries for the MiniSeq
Experiment: |  |  | Initials: | Date:
MiniSeq run information: |  |  |  | Date:
Denaturing and diluting the sample pool libraries to be loaded onto
the MiniSeq instrument
Preparation of fresh 0.1 N NaOH
1. Prepare fresh 0.1 N NaOH from 10 N NaOH stock as follows:
[_] | 1.a Add 10 μl of 10 N NaOH stock to 90 μl of nuclease-free water. Mix thoroughly.
The NaOH is now at a concentration of 1 N NaOH.
[_] | 1.b Add 50 μl of 1 N NaOH dilution from step 1.a above to 450 μl of nuclease-free water.
Mix thoroughly. The NaOH is now at a concentration of 0.1 N NaOH.
Preparation of 200 mM Tris·Cl
[_] | 1. Dilute 100 µl of the 1 M Tris·Cl, pH 7.0 stock, to 400 µl of water.
[_] | 2. Mix thoroughly. The Tris·Cl is now at a concentration of 200 mM Tris·Cl.
Denaturing and diluting the pooled library into position #16
[_] | 1. Take 5 μl of the combined pooled 4 nM lane from the Dilute Lib tab,
and add it to 5 μl of 0.1 N NaOH. The samples are now at a concentration of 2 nM.
[_] | 2. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 3. Denature the 2 nM sample for 5 min at room temperature.
[_] | 4. Add 5 μl of the 200 mM Tris·Cl, pH 7, to the 10 μl reaction from step 3.
[_] | 5. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 6. Add 985 μl of the HT1 Hybridization Buffer to the 15 μl reaction from step 5. The samples are
now at a 20 pM concentration and are ready to be diluted prior to being loaded onto the MiniSeq
Reagent Cartridge.
[_] | 7. Dilute the sample again to a final concentration appropriate to load onto the MiniSeq Reagent
Cartridge. We typically load a final concentration of 1.2 pM onto the MiniSeq Reagent Cartridge
with 10% PhiX Control Spike-in into Sample Position #16. Hence, add 467 µl of HT1 Buffer, 3 µl of a
denatured and diluted 20 pM PhiX Control to 30 µl of the 20 pM denatured library for a final volume of 500 µl.
Mix thoroughly before adding to position #16.
[_] | 8. Discard the unused 0.1 N NaOH.
MiniSeq run setup information:
| 1. Sequencing setup: | Select Category: Other > FASTQ Only
Read Type: Paired End Read
Cycles: 74 for R1, 74 for R2
Indices: 10 bp
2. Sequenced with 10% PhiX Control

## 8. Denat Dilut Lib NextSeq
How to Denature and Dilute Libraries for the NextSeq
Experiment: |  |  | Initials: | Date:
NextSeq run information: |  |  |  | Date:
Denaturing and diluting the sample pool libraries to be loaded onto the
NextSeq instrument
Preparation of fresh 0.2 N NaOH
1. Prepare fresh 0.2 N NaOH from 10 N NaOH stock as follows:
[_] | 1.a Add 20 μl of 10 N NaOH stock to 180 μl of nuclease-free water. Mix thoroughly.
The NaOH is now at a concentration of 1 N NaOH.
[_] | 1.b Add 100 μl of 1 N NaOH dilution from step 1.a above to 400 μl of nuclease-free water.
Mix thoroughly. The NaOH is now at a concentration of 0.2 N NaOH.
Preparation of 200 mM Tris·Cl
[_] | 1. Dilute 100 µl of the 1 M Tris·Cl, pH 7.0 stock, to 400 µl of water.
[_] | 2. Mix thoroughly. The Tris·Cl is now at a concentration of 200 mM Tris·Cl.
Denaturing and diluting the pooled library into position #10
[_] | 1. Take 5 μl of the combined pooled 4 nM lane from the Dilute Lib tab,
and add it to 5 μl of 0.2 N NaOH. The samples are now at a concentration of 2 nM.
[_] | 2. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 3. Denature the 2 nM sample for 5 min at room temperature.
[_] | 4. Add 5 μl of the 200 mM Tris·Cl, pH 7, to the 10 μl reaction from step 3.
[_] | 5. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 6. Add 985 μl of the HT1 Hybridization Buffer to the 15 μl reaction from step 5. The samples are
now at a 20 pM concentration and are ready to be diluted prior to being loaded onto the NextSeq
Reagent Cartridge.
[_] | 7. Dilute the sample again to a final concentration appropriate to load onto the NextSeq Reagent
Cartridge. We typically load a final concentration of 1.2 pM onto the NextSeq Reagent Cartridge
with 10% PhiX Control Spike-in into Sample Position #10. Hence, add 2 aliquots 607.1 µl (1.2142 ml) of HT1
Buffer, 7.8 µl denatured and diluted 20 pM PhiX Control to 78 µl of the 20 pM denatured library for a final
volume of 1.3 ml. Mix thoroughly before adding to position #10.
[_] | 8. Discard the unused 0.2 N NaOH.
NextSeq run setup information:
| 1. Sequencing setup: | Select Category: Other > FASTQ Only
Read Type: Paired End Read
Cycles: 74 for R1, 74 for R2
Indices: 10 bp
2. Sequenced with 10% PhiX Control

## 9. Import Sample Sheet
[Header]
IEMFileVersion | 4
Investigator Name
Experiment Name
Date
Workflow | GenerateFASTQ
Application | FASTQ Only
Assay | QIAseq RUDI
Description
Chemistry | Amplicon
[Reads]
74
74
[Settings]
ReverseComplement | 0
Adapter | AGATCGGAAGAGCACACGTCTGAACTCCAGTCA
AdapterRead2 | AGATCGGAAGAGCGTCGTGTAGGGAAAGAGTGT
[Data]
Sample_ID | Sample_Name | Sample_Plate | Sample_Well | I7_Index_ID | index | I5_Index_ID | index2 | Sample_Project | Description
0 | 0 |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A
|  |  |  | #N/A | #N/A | #N/A | #N/A

## 10. CLC DEMUX and SID-ID
ATGTCTTACG | A01
CAAGTAGTCT | B01
GGATAGGTCT | C01
TACACTACGA | D01
ATATTCGACG | E01
CAGTCCAATT | F01
GACAACTGTC | G01
TCAAGCAATC | H01
ATCGTGTTCT | A02
CATTCGCGTC | B02
GAGAGCTACT | C02
TGACAGCGTG | D02
ATGTAGGCCT | E02
CCGATGGTTG | F02
GCTTCCATTC | G02
TGGAACTCGG | H02
AAGTTCGTGC | A03
CGGTCTCTAT | B03
GAACAATCCA | C03
TTAACACGAG | D03
ATGTGGAGGT | E03
CTACCGGTTC | F03
GTACCATGCC | G03
TTAATGGAGG | H03
CTTCAATCCT | A04
GAGATAACTG | B04
TAATCTAGGC | C04
GCTTAGTCAT | D04
CTTGGTGATT | E04
TCCAGAGTCC | F04
GTAATAGCCA | G04
TGCCTAACGC | H04
TCGATTACCA | A05
GAGCAGCCTT | B05
CCGGTAAGCT | C05
TACCGTTCTT | D05
GGCATAGTGG | E05
CTTGTGTCGG | F05
TTGGCATTAC | G05
CTGGAGGATA | H05
CGGTTGTAGC | A06
GGACCTTATT | B06
TACGAGCTGC | C06
CATTGCATAG | D06
GTAGGAACTT | E06
TCGCATCACG | F06
CTTCTGACAT | G06
TTCTTCCAAC | H06
GTACACTCAT | A07
TGTCAGGCTC | B07
CATTGGTGCG | C07
GACCTTCGTC | D07
CCGTCTGAAC | E07
GCATATCCTT | F07
TGCCTGATTA | G07
GAGTAAGGAG | H07
TCGCGTGGTA | A08
CGGTTATCCG | B08
GAACTTGTTG | C08
CTACTCAGAC | D08
TCTGAGCCTT | E08
CCGCTACTTA | F08
GTTCATGGTA | G08
TACGGCAACG | H08
CGGTTCAGTG | A09
GAAGGCATCT | B09
TCGTCCGTCA | C09
CTCCTTGCCA | D09
TCACACAGTA | E09
GAAGTACAGT | F09
TGGAGTCTTG | G09
CTAGTGCTTA | H09
GCTATGAATC | A10
CTCGTCATCC | B10
TCTCGGTTCG | C10
CCGCCTTACT | D10
GCATGTTAGT | E10
CTGGTGTGCA | F10
TAGACACCAA | G10
GAAGCGTGAT | H10
GCAGTGATCG | A11
CTACCTCTCT | B11
TACGGTTAGA | C11
GTGTTACTGA | D11
TTGGTTAGTC | E11
CTATCCGCTG | F11
GCTCAGCTTG | G11
CATTGACTCT | H11
GAGGTGAACA | A12
CTATCCTAGC | B12
TGTCCTGTAC | C12
GAACACCACG | D12
CCGTCTTGTG | E12
TGCTAGTAGT | F12
GCTCATTGCT | G12
CTGTGGACAC | H12

## 11. SID Plate Layouts
SID-TS-96S RT Plate layout: SID-TS-96S Plate for 96 Samples
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | A01 | A02 | A03 | A04 | A05 | A06 | A07 | A08 | A09 | A10 | A11 | A12
B | B01 | B02 | B03 | B04 | B05 | B06 | B07 | B08 | B09 | B10 | B11 | B12
C | C01 | C02 | C03 | C04 | C05 | C06 | C07 | C08 | C09 | C10 | C11 | C12
D | D01 | D02 | D03 | D04 | D05 | D06 | D07 | D08 | D09 | D10 | D11 | D12
E | E01 | E02 | E03 | E04 | E05 | E06 | E07 | E08 | E09 | E10 | E11 | E12
F | F01 | F02 | F03 | F04 | F05 | F06 | F07 | F08 | F09 | F10 | F11 | F12
G | G01 | G02 | G03 | G04 | G05 | G06 | G07 | G08 | G09 | G10 | G11 | G12
H | H01 | H02 | H03 | H04 | H05 | H06 | H07 | H08 | H09 | H10 | H11 | H12
SID-TS-24S RT Plate layout: SID-TS-24S Plate for 24 Samples
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | A01 | A02 | A03 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
B | B01 | B02 | B03 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
C | C01 | C02 | C03 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
D | D01 | D02 | D03 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
E | E01 | E02 | E03 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
F | F01 | F02 | F03 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
G | G01 | G02 | G03 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
H | H01 | H02 | H03 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty

## 12. QIAseq UX Index Plates
QIAseq UX 96 Index Kit IL UDI Set A (96) (cat. no. 331815) layout: RUDI-96AA
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | iRUDI-001 | iRUDI-009 | iRUDI-017 | iRUDI-025 | iRUDI-033 | iRUDI-041 | iRUDI-049 | iRUDI-057 | iRUDI-065 | iRUDI-073 | iRUDI-081 | iRUDI-089
B | iRUDI-002 | iRUDI-010 | iRUDI-018 | iRUDI-026 | iRUDI-034 | iRUDI-042 | iRUDI-050 | iRUDI-058 | iRUDI-066 | iRUDI-074 | iRUDI-082 | iRUDI-090
C | iRUDI-003 | iRUDI-011 | iRUDI-019 | iRUDI-027 | iRUDI-035 | iRUDI-043 | iRUDI-051 | iRUDI-059 | iRUDI-067 | iRUDI-075 | iRUDI-083 | iRUDI-091
D | iRUDI-004 | iRUDI-012 | iRUDI-020 | iRUDI-028 | iRUDI-036 | iRUDI-044 | iRUDI-052 | iRUDI-060 | iRUDI-068 | iRUDI-076 | iRUDI-084 | iRUDI-092
E | iRUDI-005 | iRUDI-013 | iRUDI-021 | iRUDI-029 | iRUDI-037 | iRUDI-045 | iRUDI-053 | iRUDI-061 | iRUDI-069 | iRUDI-077 | iRUDI-085 | iRUDI-093
F | iRUDI-006 | iRUDI-014 | iRUDI-022 | iRUDI-030 | iRUDI-038 | iRUDI-046 | iRUDI-054 | iRUDI-062 | iRUDI-070 | iRUDI-078 | iRUDI-086 | iRUDI-094
G | iRUDI-007 | iRUDI-015 | iRUDI-023 | iRUDI-031 | iRUDI-039 | iRUDI-047 | iRUDI-055 | iRUDI-063 | iRUDI-071 | iRUDI-079 | iRUDI-087 | iRUDI-095
H | iRUDI-008 | iRUDI-016 | iRUDI-024 | iRUDI-032 | iRUDI-040 | iRUDI-048 | iRUDI-056 | iRUDI-064 | iRUDI-072 | iRUDI-080 | iRUDI-088 | iRUDI-096
* Unique i5 and i7 Index pairs are wet single-use plates.
QIAseq UX 96 Index Kit IL UDI Set B (96) (cat. no. 331825) layout: RUDI-96BA
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | iRUDI-097 | iRUDI-105 | iRUDI-113 | iRUDI-121 | iRUDI-129 | iRUDI-137 | iRUDI-145 | iRUDI-153 | iRUDI-161 | iRUDI-169 | iRUDI-177 | iRUDI-185
B | iRUDI-098 | iRUDI-106 | iRUDI-114 | iRUDI-122 | iRUDI-130 | iRUDI-138 | iRUDI-146 | iRUDI-154 | iRUDI-162 | iRUDI-170 | iRUDI-178 | iRUDI-186
C | iRUDI-099 | iRUDI-107 | iRUDI-115 | iRUDI-123 | iRUDI-131 | iRUDI-139 | iRUDI-147 | iRUDI-155 | iRUDI-163 | iRUDI-171 | iRUDI-179 | iRUDI-187
D | iRUDI-100 | iRUDI-108 | iRUDI-116 | iRUDI-124 | iRUDI-132 | iRUDI-140 | iRUDI-148 | iRUDI-156 | iRUDI-164 | iRUDI-172 | iRUDI-180 | iRUDI-188
E | iRUDI-101 | iRUDI-109 | iRUDI-117 | iRUDI-125 | iRUDI-133 | iRUDI-141 | iRUDI-149 | iRUDI-157 | iRUDI-165 | iRUDI-173 | iRUDI-181 | iRUDI-189
F | iRUDI-102 | iRUDI-110 | iRUDI-118 | iRUDI-126 | iRUDI-134 | iRUDI-142 | iRUDI-150 | iRUDI-158 | iRUDI-166 | iRUDI-174 | iRUDI-182 | iRUDI-190
G | iRUDI-103 | iRUDI-111 | iRUDI-119 | iRUDI-127 | iRUDI-135 | iRUDI-143 | iRUDI-151 | iRUDI-159 | iRUDI-167 | iRUDI-175 | iRUDI-183 | iRUDI-191
H | iRUDI-104 | iRUDI-112 | iRUDI-120 | iRUDI-128 | iRUDI-136 | iRUDI-144 | iRUDI-152 | iRUDI-160 | iRUDI-168 | iRUDI-176 | iRUDI-184 | iRUDI-192
* Unique i5 and i7 Index pairs are wet single-use plates.
QIAseq UX 96 Index Kit IL UDI Set C (96) (cat. no. 331835) layout: RUDI-96CA
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | iRUDI-193 | iRUDI-201 | iRUDI-209 | iRUDI-217 | iRUDI-225 | iRUDI-233 | iRUDI-241 | iRUDI-249 | iRUDI-257 | iRUDI-265 | iRUDI-273 | iRUDI-281
B | iRUDI-194 | iRUDI-202 | iRUDI-210 | iRUDI-218 | iRUDI-226 | iRUDI-234 | iRUDI-242 | iRUDI-250 | iRUDI-258 | iRUDI-266 | iRUDI-274 | iRUDI-282
C | iRUDI-195 | iRUDI-203 | iRUDI-211 | iRUDI-219 | iRUDI-227 | iRUDI-235 | iRUDI-243 | iRUDI-251 | iRUDI-259 | iRUDI-267 | iRUDI-275 | iRUDI-283
D | iRUDI-196 | iRUDI-204 | iRUDI-212 | iRUDI-220 | iRUDI-228 | iRUDI-236 | iRUDI-244 | iRUDI-252 | iRUDI-260 | iRUDI-268 | iRUDI-276 | iRUDI-284
E | iRUDI-197 | iRUDI-205 | iRUDI-213 | iRUDI-221 | iRUDI-229 | iRUDI-237 | iRUDI-245 | iRUDI-253 | iRUDI-261 | iRUDI-269 | iRUDI-277 | iRUDI-285
F | iRUDI-198 | iRUDI-206 | iRUDI-214 | iRUDI-222 | iRUDI-230 | iRUDI-238 | iRUDI-246 | iRUDI-254 | iRUDI-262 | iRUDI-270 | iRUDI-278 | iRUDI-286
G | iRUDI-199 | iRUDI-207 | iRUDI-215 | iRUDI-223 | iRUDI-231 | iRUDI-239 | iRUDI-247 | iRUDI-255 | iRUDI-263 | iRUDI-271 | iRUDI-279 | iRUDI-287
H | iRUDI-200 | iRUDI-208 | iRUDI-216 | iRUDI-224 | iRUDI-232 | iRUDI-240 | iRUDI-248 | iRUDI-256 | iRUDI-264 | iRUDI-272 | iRUDI-280 | iRUDI-288
* Unique i5 and i7 Index pairs are wet single-use plates.
QIAseq UX 96 Index Kit IL UDI Set D (96) (cat. no. 331845) layout: RUDI-96DA
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | iRUDI-289 | iRUDI-297 | iRUDI-305 | iRUDI-313 | iRUDI-321 | iRUDI-329 | iRUDI-337 | iRUDI-345 | iRUDI-353 | iRUDI-361 | iRUDI-369 | iRUDI-377
B | iRUDI-290 | iRUDI-298 | iRUDI-306 | iRUDI-314 | iRUDI-322 | iRUDI-330 | iRUDI-338 | iRUDI-346 | iRUDI-354 | iRUDI-362 | iRUDI-370 | iRUDI-378
C | iRUDI-291 | iRUDI-299 | iRUDI-307 | iRUDI-315 | iRUDI-323 | iRUDI-331 | iRUDI-339 | iRUDI-347 | iRUDI-355 | iRUDI-363 | iRUDI-371 | iRUDI-379
D | iRUDI-292 | iRUDI-300 | iRUDI-308 | iRUDI-316 | iRUDI-324 | iRUDI-332 | iRUDI-340 | iRUDI-348 | iRUDI-356 | iRUDI-364 | iRUDI-372 | iRUDI-380
E | iRUDI-293 | iRUDI-301 | iRUDI-309 | iRUDI-317 | iRUDI-325 | iRUDI-333 | iRUDI-341 | iRUDI-349 | iRUDI-357 | iRUDI-365 | iRUDI-373 | iRUDI-381
F | iRUDI-294 | iRUDI-302 | iRUDI-310 | iRUDI-318 | iRUDI-326 | iRUDI-334 | iRUDI-342 | iRUDI-350 | iRUDI-358 | iRUDI-366 | iRUDI-374 | iRUDI-382
G | iRUDI-295 | iRUDI-303 | iRUDI-311 | iRUDI-319 | iRUDI-327 | iRUDI-335 | iRUDI-343 | iRUDI-351 | iRUDI-359 | iRUDI-367 | iRUDI-375 | iRUDI-383
H | iRUDI-296 | iRUDI-304 | iRUDI-312 | iRUDI-320 | iRUDI-328 | iRUDI-336 | iRUDI-344 | iRUDI-352 | iRUDI-360 | iRUDI-368 | iRUDI-376 | iRUDI-384
* Unique i5 and i7 Index pairs are wet single-use plates.
QIAseq UX 96 Index Kit IL UDI Set E (96) (cat. no. 331855) layout: RUDI-96EA
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | iRUDI-385 | iRUDI-393 | iRUDI-401 | iRUDI-409 | iRUDI-417 | iRUDI-425 | iRUDI-433 | iRUDI-441 | iRUDI-449 | iRUDI-457 | iRUDI-465 | iRUDI-473
B | iRUDI-386 | iRUDI-394 | iRUDI-402 | iRUDI-410 | iRUDI-418 | iRUDI-426 | iRUDI-434 | iRUDI-442 | iRUDI-450 | iRUDI-458 | iRUDI-466 | iRUDI-474
C | iRUDI-387 | iRUDI-395 | iRUDI-403 | iRUDI-411 | iRUDI-419 | iRUDI-427 | iRUDI-435 | iRUDI-443 | iRUDI-451 | iRUDI-459 | iRUDI-467 | iRUDI-475
D | iRUDI-388 | iRUDI-396 | iRUDI-404 | iRUDI-412 | iRUDI-420 | iRUDI-428 | iRUDI-436 | iRUDI-444 | iRUDI-452 | iRUDI-460 | iRUDI-468 | iRUDI-476
E | iRUDI-389 | iRUDI-397 | iRUDI-405 | iRUDI-413 | iRUDI-421 | iRUDI-429 | iRUDI-437 | iRUDI-445 | iRUDI-453 | iRUDI-461 | iRUDI-469 | iRUDI-477
F | iRUDI-390 | iRUDI-398 | iRUDI-406 | iRUDI-414 | iRUDI-422 | iRUDI-430 | iRUDI-438 | iRUDI-446 | iRUDI-454 | iRUDI-462 | iRUDI-470 | iRUDI-478
G | iRUDI-391 | iRUDI-399 | iRUDI-407 | iRUDI-415 | iRUDI-423 | iRUDI-431 | iRUDI-439 | iRUDI-447 | iRUDI-455 | iRUDI-463 | iRUDI-471 | iRUDI-479
H | iRUDI-392 | iRUDI-400 | iRUDI-408 | iRUDI-416 | iRUDI-424 | iRUDI-432 | iRUDI-440 | iRUDI-448 | iRUDI-456 | iRUDI-464 | iRUDI-472 | iRUDI-480
* Unique i5 and i7 Index pairs are wet single-use plates.
QIAseq UX 96 Index Kit IL UDI Set F (96) (cat. no. 331865) layout: RUDI-96FA
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | iRUDI-481 | iRUDI-489 | iRUDI-497 | iRUDI-505 | iRUDI-513 | iRUDI-521 | iRUDI-529 | iRUDI-537 | iRUDI-545 | iRUDI-553 | iRUDI-561 | iRUDI-569
B | iRUDI-482 | iRUDI-490 | iRUDI-498 | iRUDI-506 | iRUDI-514 | iRUDI-522 | iRUDI-530 | iRUDI-538 | iRUDI-546 | iRUDI-554 | iRUDI-562 | iRUDI-570
C | iRUDI-483 | iRUDI-491 | iRUDI-499 | iRUDI-507 | iRUDI-515 | iRUDI-523 | iRUDI-531 | iRUDI-539 | iRUDI-547 | iRUDI-555 | iRUDI-563 | iRUDI-571
D | iRUDI-484 | iRUDI-492 | iRUDI-500 | iRUDI-508 | iRUDI-516 | iRUDI-524 | iRUDI-532 | iRUDI-540 | iRUDI-548 | iRUDI-556 | iRUDI-564 | iRUDI-572
E | iRUDI-485 | iRUDI-493 | iRUDI-501 | iRUDI-509 | iRUDI-517 | iRUDI-525 | iRUDI-533 | iRUDI-541 | iRUDI-549 | iRUDI-557 | iRUDI-565 | iRUDI-573
F | iRUDI-486 | iRUDI-494 | iRUDI-502 | iRUDI-510 | iRUDI-518 | iRUDI-526 | iRUDI-534 | iRUDI-542 | iRUDI-550 | iRUDI-558 | iRUDI-566 | iRUDI-574
G | iRUDI-487 | iRUDI-495 | iRUDI-503 | iRUDI-511 | iRUDI-519 | iRUDI-527 | iRUDI-535 | iRUDI-543 | iRUDI-551 | iRUDI-559 | iRUDI-567 | iRUDI-575
H | iRUDI-488 | iRUDI-496 | iRUDI-504 | iRUDI-512 | iRUDI-520 | iRUDI-528 | iRUDI-536 | iRUDI-544 | iRUDI-552 | iRUDI-560 | iRUDI-568 | iRUDI-576
* Unique i5 and i7 Index pairs are wet single-use plates.
QIAseq UX 96 Index Kit IL UDI Set G (96) (cat. no. 331875) layout: RUDI-96GA
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | iRUDI-577 | iRUDI-585 | iRUDI-593 | iRUDI-601 | iRUDI-609 | iRUDI-617 | iRUDI-625 | iRUDI-633 | iRUDI-641 | iRUDI-649 | iRUDI-657 | iRUDI-665
B | iRUDI-578 | iRUDI-586 | iRUDI-594 | iRUDI-602 | iRUDI-610 | iRUDI-618 | iRUDI-626 | iRUDI-634 | iRUDI-642 | iRUDI-650 | iRUDI-658 | iRUDI-666
C | iRUDI-579 | iRUDI-587 | iRUDI-595 | iRUDI-603 | iRUDI-611 | iRUDI-619 | iRUDI-627 | iRUDI-635 | iRUDI-643 | iRUDI-651 | iRUDI-659 | iRUDI-667
D | iRUDI-580 | iRUDI-588 | iRUDI-596 | iRUDI-604 | iRUDI-612 | iRUDI-620 | iRUDI-628 | iRUDI-636 | iRUDI-644 | iRUDI-652 | iRUDI-660 | iRUDI-668
E | iRUDI-581 | iRUDI-589 | iRUDI-597 | iRUDI-605 | iRUDI-613 | iRUDI-621 | iRUDI-629 | iRUDI-637 | iRUDI-645 | iRUDI-653 | iRUDI-661 | iRUDI-669
F | iRUDI-582 | iRUDI-590 | iRUDI-598 | iRUDI-606 | iRUDI-614 | iRUDI-622 | iRUDI-630 | iRUDI-638 | iRUDI-646 | iRUDI-654 | iRUDI-662 | iRUDI-670
G | iRUDI-583 | iRUDI-591 | iRUDI-599 | iRUDI-607 | iRUDI-615 | iRUDI-623 | iRUDI-631 | iRUDI-639 | iRUDI-647 | iRUDI-655 | iRUDI-663 | iRUDI-671
H | iRUDI-584 | iRUDI-592 | iRUDI-600 | iRUDI-608 | iRUDI-616 | iRUDI-624 | iRUDI-632 | iRUDI-640 | iRUDI-648 | iRUDI-656 | iRUDI-664 | iRUDI-672
* Unique i5 and i7 Index pairs are wet single-use plates.
QIAseq UX 96 Index Kit IL UDI Set H (96) (cat. no. 331885) layout: RUDI-96HA
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | iRUDI-673 | iRUDI-681 | iRUDI-689 | iRUDI-697 | iRUDI-705 | iRUDI-713 | iRUDI-721 | iRUDI-729 | iRUDI-737 | iRUDI-745 | iRUDI-753 | iRUDI-761
B | iRUDI-674 | iRUDI-682 | iRUDI-690 | iRUDI-698 | iRUDI-706 | iRUDI-714 | iRUDI-722 | iRUDI-730 | iRUDI-738 | iRUDI-746 | iRUDI-754 | iRUDI-762
C | iRUDI-675 | iRUDI-683 | iRUDI-691 | iRUDI-699 | iRUDI-707 | iRUDI-715 | iRUDI-723 | iRUDI-731 | iRUDI-739 | iRUDI-747 | iRUDI-755 | iRUDI-763
D | iRUDI-676 | iRUDI-684 | iRUDI-692 | iRUDI-700 | iRUDI-708 | iRUDI-716 | iRUDI-724 | iRUDI-732 | iRUDI-740 | iRUDI-748 | iRUDI-756 | iRUDI-764
E | iRUDI-677 | iRUDI-685 | iRUDI-693 | iRUDI-701 | iRUDI-709 | iRUDI-717 | iRUDI-725 | iRUDI-733 | iRUDI-741 | iRUDI-749 | iRUDI-757 | iRUDI-765
F | iRUDI-678 | iRUDI-686 | iRUDI-694 | iRUDI-702 | iRUDI-710 | iRUDI-718 | iRUDI-726 | iRUDI-734 | iRUDI-742 | iRUDI-750 | iRUDI-758 | iRUDI-766
G | iRUDI-679 | iRUDI-687 | iRUDI-695 | iRUDI-703 | iRUDI-711 | iRUDI-719 | iRUDI-727 | iRUDI-735 | iRUDI-743 | iRUDI-751 | iRUDI-759 | iRUDI-767
H | iRUDI-680 | iRUDI-688 | iRUDI-696 | iRUDI-704 | iRUDI-712 | iRUDI-720 | iRUDI-728 | iRUDI-736 | iRUDI-744 | iRUDI-752 | iRUDI-760 | iRUDI-768
* Unique i5 and i7 Index pairs are wet single-use plates.
QIAseq UX 12 Index Kit IL UDI (12) (cat. no. 331801) layout: RUDI-12A
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | iRUDI-001 | iRUDI-009 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
B | iRUDI-002 | iRUDI-010 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
C | iRUDI-003 | iRUDI-011 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
D | iRUDI-004 | iRUDI-012 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
E | iRUDI-005 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
F | iRUDI-006 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
G | iRUDI-007 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
H | iRUDI-008 | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty | Empty
* Unique i5 and i7 Index pairs are wet single-use plates.

## 13. UPXome library structure
Library structure based on the chosen RT primers
N6-T RT Primer + ODT-T RT Primer | N6-T RT Primer | (P5 side)-CTACACGACGCTCTTCCGATCT--[Insert]--[CCCTGC]--[10bpID]—AGATCGGAAGAGCACACGTCTG-(P7 side)
| ODT-T RT Primer | (P5 side)-CTACACGACGCTCTTCCGATCT--TTTVN[Insert]--[CCCTGC]--[10bpID]—AGATCGGAAGAGCACACGTCTG-(P7 side)
N6-T RT Primer |  | (P5 side)-CTACACGACGCTCTTCCGATCT--[Insert]--[CCCTGC]--[10bpID]—AGATCGGAAGAGCACACGTCTG-(P7 side)
ODT-T RT Primer |  | (P5 side)-CTACACGACGCTCTTCCGATCT--TTTVN[Insert]--[CCCTGC]--[10bpID]—AGATCGGAAGAGCACACGTCTG-(P7 side)

## Autopopulate UDI's (hidden)
| i7 Index Name | i7 Bases for Entry on Sample Sheet |  | i5 Index Name | i5 Bases for Entry on DNA Sample Sheet | i5 Bases Reverse Complement for Entry on NovaSeq Sample Sheet
PlateA_A01 | i7MUDI-001 | TGAACGTTGT | PlateA_A01 | i5MUDI-001 | ATGGCCGACT | AGTCGGCCAT
PlateA_B01 | i7MUDI-002 | ACCAGACTTG | PlateA_B01 | i5MUDI-002 | CGATGAGCAC | GTGCTCATCG
PlateA_C01 | i7MUDI-003 | ACTGGCGAAC | PlateA_C01 | i5MUDI-003 | GATAAGTCGA | TCGACTTATC
PlateA_D01 | i7MUDI-004 | GCGTTAGGCA | PlateA_D01 | i5MUDI-004 | TCACGCCTTG | CAAGGCGTGA
PlateA_E01 | i7MUDI-005 | TTATCGGCCT | PlateA_E01 | i5MUDI-005 | AGGAACACAA | TTGTGTTCCT
PlateA_F01 | i7MUDI-006 | GAGGTATAAG | PlateA_F01 | i5MUDI-006 | CTCAGTAGGC | GCCTACTGAG
PlateA_G01 | i7MUDI-007 | TCAAGGATTC | PlateA_G01 | i5MUDI-007 | GAAGTGCCTG | CAGGCACTTC
PlateA_H01 | i7MUDI-008 | CGAACCGAGA | PlateA_H01 | i5MUDI-008 | TCTCTCGCCT | AGGCGAGAGA
PlateA_A02 | i7MUDI-009 | GAGCCAAGTT | PlateA_A02 | i5MUDI-009 | AGGCACCTTC | GAAGGTGCCT
PlateA_B02 | i7MUDI-010 | AAGGCCGTAG | PlateA_B02 | i5MUDI-010 | CTGTTGGTAA | TTACCAACAG
PlateA_C02 | i7MUDI-011 | TTAGAGAAGC | PlateA_C02 | i5MUDI-011 | GCTGGTACCT | AGGTACCAGC
PlateA_D02 | i7MUDI-012 | TCTAAGACCA | PlateA_D02 | i5MUDI-012 | TAAGGAGCGG | CCGCTCCTTA
PlateA_E02 | i7MUDI-013 | TGTAACCACT | PlateA_E02 | i5MUDI-013 | AATCGCTCCA | TGGAGCGATT
PlateA_F02 | i7MUDI-014 | CCGACACAAG | PlateA_F02 | i5MUDI-014 | CTCCTAATTG | CAATTAGGAG
PlateA_G02 | i7MUDI-015 | CTCTGATGGC | PlateA_G02 | i5MUDI-015 | GCCTCATAAT | ATTATGAGGC
PlateA_H02 | i7MUDI-016 | CGGCCTGTTA | PlateA_H02 | i5MUDI-016 | TGTATTGAGC | GCTCAATACA
PlateA_A03 | i7MUDI-017 | TGCATAGCTT | PlateA_A03 | i5MUDI-017 | AGCCATAACA | TGTTATGGCT
PlateA_B03 | i7MUDI-018 | AACCTTCTCG | PlateA_B03 | i5MUDI-018 | CCACAAGTGG | CCACTTGTGG
PlateA_C03 | i7MUDI-019 | AAGAGATCAC | PlateA_C03 | i5MUDI-019 | GTTATCACAC | GTGTGATAAC
PlateA_D03 | i7MUDI-020 | GCCTGAAGGA | PlateA_D03 | i5MUDI-020 | TACCGTTCTT | AAGAACGGTA
PlateA_E03 | i7MUDI-021 | ATTGTGCCTT | PlateA_E03 | i5MUDI-021 | AGGCGTTAGG | CCTAACGCCT
PlateA_F03 | i7MUDI-769 | CTCCTTGGCA | PlateA_F03 | i5MUDI-769 | TGGTTGGCGA | TCGCCAACCA
PlateA_G03 | i7MUDI-023 | TACCATGAAC | PlateA_G03 | i5MUDI-023 | GTAATAGCCA | TGGCTATTAC
PlateA_H03 | i7MUDI-024 | CATTGGCAGA | PlateA_H03 | i5MUDI-024 | TAGCGCCGAT | ATCGGCGCTA
PlateA_A04 | i7MUDI-025 | CACTGCTATT | PlateA_A04 | i5MUDI-025 | CATTCTTGGA | TCCAAGAATG
PlateA_B04 | i7MUDI-026 | AATGGTAGGT | PlateA_B04 | i5MUDI-026 | ATGCAAGGTT | AACCTTGCAT
PlateA_C04 | i7MUDI-027 | GATACCTATG | PlateA_C04 | i5MUDI-027 | CGCCAGACAA | TTGTCTGGCG
PlateA_D04 | i7MUDI-028 | CACTAGGTAC | PlateA_D04 | i5MUDI-028 | GAAGGTTGGC | GCCAACCTTC
PlateA_E04 | i7MUDI-029 | AGCTCGTTCA | PlateA_E04 | i5MUDI-029 | TCGCATCACG | CGTGATGCGA
PlateA_F04 | i7MUDI-030 | TGTCAGTCTT | PlateA_F04 | i5MUDI-030 | CCGGTCATGA | TCATGACCGG
PlateA_G04 | i7MUDI-031 | GATGAACAGT | PlateA_G04 | i5MUDI-031 | ATTCACAAGC | GCTTGTGAAT
PlateA_H04 | i7MUDI-032 | ACAATCGGCG | PlateA_H04 | i5MUDI-032 | CAACCTGTAA | TTACAGGTTG
PlateA_A05 | i7MUDI-033 | GATTGAGTTC | PlateA_A05 | i5MUDI-033 | GCCAGTCGTT | AACGACTGGC
PlateA_B05 | i7MUDI-034 | GTAATGCCAA | PlateA_B05 | i5MUDI-034 | TGCCTTGTCG | CGACAAGGCA
PlateA_C05 | i7MUDI-035 | TCGTTGCGCT | PlateA_C05 | i5MUDI-035 | CTATCCGCTG | CAGCGGATAG
PlateA_D05 | i7MUDI-036 | AGGTGAGTAT | PlateA_D05 | i5MUDI-036 | AATGCCGGAA | TTCCGGCATT
PlateA_E05 | i7MUDI-037 | TCGATAATGG | PlateA_E05 | i5MUDI-037 | CGGTTATCCG | CGGATAACCG
PlateA_F05 | i7MUDI-038 | GCGTCTCTTC | PlateA_F05 | i5MUDI-038 | GCGGAAGAGT | ACTCTTCCGC
PlateA_G05 | i7MUDI-039 | GTCTCCTGCA | PlateA_G05 | i5MUDI-039 | TTGGTTAGTC | GACTAACCAA
PlateA_H05 | i7MUDI-040 | GAGCTTCATT | PlateA_H05 | i5MUDI-040 | TTCAGTGTGA | TCACACTGAA
PlateA_A06 | i7MUDI-041 | AGGCCTACAT | PlateA_A06 | i5MUDI-041 | AGAATTCTGG | CCAGAATTCT
PlateA_B06 | i7MUDI-042 | TGTGGAACCG | PlateA_B06 | i5MUDI-042 | CATTGACTCT | AGAGTCAATG
PlateA_C06 | i7MUDI-043 | CGTATTAAGC | PlateA_C06 | i5MUDI-043 | GCGGCTTCAA | TTGAAGCCGC
PlateA_D06 | i7MUDI-044 | CCAGTGGTTA | PlateA_D06 | i5MUDI-044 | TTATGGTCTC | GAGACCATAA
PlateA_E06 | i7MUDI-045 | GCGTTCGAGT | PlateA_E06 | i5MUDI-045 | CGTAACCAGG | CCTGGTTACG
PlateA_F06 | i7MUDI-046 | CCTTCCGGTT | PlateA_F06 | i5MUDI-046 | AGCTCAGATA | TATCTGAGCT
PlateA_G06 | i7MUDI-047 | CACAAGACGG | PlateA_G06 | i5MUDI-047 | CCGGTGTTAC | GTAACACCGG
PlateA_H06 | i7MUDI-048 | GCTTACACAC | PlateA_H06 | i5MUDI-048 | GACCTAACCT | AGGTTAGGTC
PlateA_A07 | i7MUDI-049 | AGGATGTCCA | PlateA_A07 | i5MUDI-049 | TTGTAGAAGG | CCTTCTACAA
PlateA_B07 | i7MUDI-050 | CACCTTATGT | PlateA_B07 | i5MUDI-050 | CCTAGCACTA | TAGTGCTAGG
PlateA_C07 | i7MUDI-051 | AAGCGGCTGT | PlateA_C07 | i5MUDI-051 | ATCGTGTTCT | AGAACACGAT
PlateA_D07 | i7MUDI-052 | TTCCTGTGAG | PlateA_D07 | i5MUDI-052 | CCAACTTATC | GATAAGTTGG
PlateA_E07 | i7MUDI-053 | AGTACAGTTC | PlateA_E07 | i5MUDI-053 | GAAGCCAAGG | CCTTGGCTTC
PlateA_F07 | i7MUDI-054 | TACAGCCTCA | PlateA_F07 | i5MUDI-054 | TGGAGTTCAA | TTGAACTCCA
PlateA_G07 | i7MUDI-055 | GTTCTATTGG | PlateA_G07 | i5MUDI-055 | CTTCAATCCT | AGGATTGAAG
PlateA_H07 | i7MUDI-056 | ATATACCGGT | PlateA_H07 | i5MUDI-056 | ATCTTGCGTG | CACGCAAGAT
PlateA_A08 | i7MUDI-057 | CCTCGGAATG | PlateA_A08 | i5MUDI-057 | CGTCTAAGGT | ACCTTAGACG
PlateA_B08 | i7MUDI-058 | GTTCTGGAAC | PlateA_B08 | i5MUDI-058 | GAGGTGAACA | TGTTCACCTC
PlateA_C08 | i7MUDI-059 | AGATTCACCA | PlateA_C08 | i5MUDI-059 | TCAGAACTAC | GTAGTTCTGA
PlateA_D08 | i7MUDI-060 | TCGGTCAGAT | PlateA_D08 | i5MUDI-060 | CGGATATTGA | TCAATATCCG
PlateA_E08 | i7MUDI-061 | CACTCTCGCT | PlateA_E08 | i5MUDI-061 | AGGAGTAGAT | ATCTACTCCT
PlateA_F08 | i7MUDI-062 | GTTGGTCCAG | PlateA_F08 | i5MUDI-062 | CCGCCGAATA | TATTCGGCGG
PlateA_G08 | i7MUDI-063 | AGCTCGAAGC | PlateA_G08 | i5MUDI-063 | GAGTCTATAC | GTATAGACTC
PlateA_H08 | i7MUDI-064 | AGAGGTTCTA | PlateA_H08 | i5MUDI-064 | TTATTACCGG | CCGGTAATAA
PlateA_A09 | i7MUDI-065 | ATGACTCGAA | PlateA_A09 | i5MUDI-065 | CGCTCGTTAG | CTAACGAGCG
PlateA_B09 | i7MUDI-066 | GAACAATCCT | PlateA_B09 | i5MUDI-066 | AACAACGCTG | CAGCGTTGTT
PlateA_C09 | i7MUDI-067 | TGGCAAGGAG | PlateA_C09 | i5MUDI-067 | CGCGGCTATT | AATAGCCGCG
PlateA_D09 | i7MUDI-068 | GAATATTGGC | PlateA_D09 | i5MUDI-068 | GCTCGACACA | TGTGTCGAGC
PlateA_E09 | i7MUDI-069 | CCGGAACCTA | PlateA_E09 | i5MUDI-069 | TTCTTCCAAC | GTTGGAAGAA
PlateA_F09 | i7MUDI-070 | ACTTGTTCGG | PlateA_F09 | i5MUDI-070 | TTGGCGGTTG | CAACCGCCAA
PlateA_G09 | i7MUDI-071 | CAAGTCCAAT | PlateA_G09 | i5MUDI-071 | AACAGGCAAT | ATTGCCTGTT
PlateA_H09 | i7MUDI-072 | AACCGCAAGG | PlateA_H09 | i5MUDI-072 | CAGAATGGCG | CGCCATTCTG
PlateA_A10 | i7MUDI-073 | ACGTTGACTC | PlateA_A10 | i5MUDI-073 | GTTGAGATTC | GAATCTCAAC
PlateA_B10 | i7MUDI-074 | CCACTTAACA | PlateA_B10 | i5MUDI-074 | TGTGTGCGGA | TCCGCACACA
PlateA_C10 | i7MUDI-075 | AGCAGTTCCT | PlateA_C10 | i5MUDI-075 | GTTCGGCGAA | TTCGCCGAAC
PlateA_D10 | i7MUDI-076 | TCGCCTTCGT | PlateA_D10 | i5MUDI-076 | AGCTGTATTG | CAATACAGCT
PlateA_E10 | i7MUDI-077 | TAGGACTGCG | PlateA_E10 | i5MUDI-077 | CAGCGGATGA | TCATCCGCTG
PlateA_F10 | i7MUDI-078 | TCCGAGCGAA | PlateA_F10 | i5MUDI-078 | GTCCTTGGAT | ATCCAAGGAC
PlateA_G10 | i7MUDI-079 | TTCGGTTGTT | PlateA_G10 | i5MUDI-079 | TCTAGATGCT | AGCATCTAGA
PlateA_H10 | i7MUDI-080 | ACAGGAGGAA | PlateA_H10 | i5MUDI-080 | CGAGCCACAT | ATGTGGCTCG
PlateA_A11 | i7MUDI-081 | CCTCCATTAA | PlateA_A11 | i5MUDI-081 | ATGGAATGGA | TCCATTCCAT
PlateA_B11 | i7MUDI-082 | AGTCGCGGTT | PlateA_B11 | i5MUDI-082 | CATTCCTCAC | GTGAGGAATG
PlateA_C11 | i7MUDI-083 | CTCATCCAGG | PlateA_C11 | i5MUDI-083 | GCATAGGAAG | CTTCCTATGC
PlateA_D11 | i7MUDI-084 | TGTGGTTGAA | PlateA_D11 | i5MUDI-084 | TGTTCGTGTT | AACACGAACA
PlateA_E11 | i7MUDI-085 | TTATGCGTGG | PlateA_E11 | i5MUDI-085 | TAAGACCGTT | AACGGTCTTA
PlateA_F11 | i7MUDI-086 | GCGAATGTAT | PlateA_F11 | i5MUDI-086 | ATGGTACCAG | CTGGTACCAT
PlateA_G11 | i7MUDI-087 | GTCAAGCTCG | PlateA_G11 | i5MUDI-087 | CCGACAGCTT | AAGCTGTCGG
PlateA_H11 | i7MUDI-088 | TAGAGTTGGA | PlateA_H11 | i5MUDI-088 | GACGATATGA | TCATATCGTC
PlateA_A12 | i7MUDI-089 | CTGATGATCT | PlateA_A12 | i5MUDI-089 | TTGTACTCCA | TGGAGTACAA
PlateA_B12 | i7MUDI-090 | ACTAGGTGTT | PlateA_B12 | i5MUDI-090 | GTGCACATAA | TTATGTGCAC
PlateA_C12 | i7MUDI-091 | CTGTTAGCGG | PlateA_C12 | i5MUDI-091 | AGGACAAGTA | TACTTGTCCT
PlateA_D12 | i7MUDI-092 | ATCGCACCAA | PlateA_D12 | i5MUDI-092 | CCGATTCGAG | CTCGAATCGG
PlateA_E12 | i7MUDI-093 | CTTACTTGGT | PlateA_E12 | i5MUDI-093 | GTAGGAACTT | AAGTTCCTAC
PlateA_F12 | i7MUDI-094 | CCTTAATGCG | PlateA_F12 | i5MUDI-094 | TACACTACGA | TCGTAGTGTA
PlateA_G12 | i7MUDI-095 | TCTCGCCTAG | PlateA_G12 | i5MUDI-095 | ATGACCTTGA | TCAAGGTCAT
PlateA_H12 | i7MUDI-096 | TCTTCAGAGA | PlateA_H12 | i5MUDI-096 | CTACGTGACG | CGTCACGTAG
PlateB_A01 | i7MUDI-097 | TACCGGTGGT | PlateB_A01 | i5MUDI-097 | AACAATCAGG | CCTGATTGTT
PlateB_B01 | i7MUDI-098 | AGGTGTTACG | PlateB_B01 | i5MUDI-098 | CTGGTGTGCA | TGCACACCAG
PlateB_C01 | i7MUDI-099 | ACAGACCGAC | PlateB_C01 | i5MUDI-099 | GCATATCCTT | AAGGATATGC
PlateB_D01 | i7MUDI-100 | CGAATACGTA | PlateB_D01 | i5MUDI-100 | TGTCCTGTAC | GTACAGGACA
PlateB_E01 | i7MUDI-101 | TAGCATCGAT | PlateB_E01 | i5MUDI-101 | AGAACGTCGC | GCGACGTTCT
PlateB_F01 | i7MUDI-102 | CCATGAGTCG | PlateB_F01 | i5MUDI-102 | CACGGACTAG | CTAGTCCGTG
PlateB_G01 | i7MUDI-103 | ACTAACATGC | PlateB_G01 | i5MUDI-103 | GTTGAACACT | AGTGTTCAAC
PlateB_H01 | i7MUDI-104 | ACACTCTCTA | PlateB_H01 | i5MUDI-104 | TCGCGTGGTA | TACCACGCGA
PlateB_A02 | i7MUDI-105 | GCTCTTGCCT | PlateB_A02 | i5MUDI-105 | AGCCACTATG | CATAGTGGCT
PlateB_B02 | i7MUDI-106 | AATCTTGAGG | PlateB_B02 | i5MUDI-106 | CCACCTACCA | TGGTAGGTGG
PlateB_C02 | i7MUDI-107 | CTTAACGGTC | PlateB_C02 | i5MUDI-107 | GTTCCGGTGT | ACACCGGAAC
PlateB_D02 | i7MUDI-108 | TTGTGACCAA | PlateB_D02 | i5MUDI-108 | TAGGTCTGAC | GTCAGACCTA
PlateB_E02 | i7MUDI-109 | TCACACACCT | PlateB_E02 | i5MUDI-109 | AGGAAGCATT | AATGCTTCCT
PlateB_F02 | i7MUDI-110 | CTGCAATTAG | PlateB_F02 | i5MUDI-110 | CCTTAGTTGG | CCAACTAAGG
PlateB_G02 | i7MUDI-111 | CTCCTTACTC | PlateB_G02 | i5MUDI-111 | GTCCTATTCA | TGAATAGGAC
PlateB_H02 | i7MUDI-112 | GCAACGCAGA | PlateB_H02 | i5MUDI-112 | TAAGATGGAC | GTCCATCTTA
PlateB_A03 | i7MUDI-113 | CCTTACCAAT | PlateB_A03 | i5MUDI-113 | AGGCCATGGT | ACCATGGCCT
PlateB_B03 | i7MUDI-114 | TTAATCCTCG | PlateB_B03 | i5MUDI-114 | CATTGGCCAA | TTGGCCAATG
PlateB_C03 | i7MUDI-115 | TTCCGAGTTC | PlateB_C03 | i5MUDI-115 | GCTATGAATC | GATTCATAGC
PlateB_D03 | i7MUDI-116 | CTCGAGAGGA | PlateB_D03 | i5MUDI-116 | TTGGTCCTCG | CGAGGACCAA
PlateB_E03 | i7MUDI-117 | TGTTGGCTGT | PlateB_E03 | i5MUDI-117 | AGCGACATAC | GTATGTCGCT
PlateB_F03 | i7MUDI-118 | CGTATCTGCG | PlateB_F03 | i5MUDI-118 | CAAGTAGTCT | AGACTACTTG
PlateB_G03 | i7MUDI-119 | CCATAGTATC | PlateB_G03 | i5MUDI-119 | GTCAAGAAGA | TCTTCTTGAC
PlateB_H03 | i7MUDI-120 | TGGACAGTAA | PlateB_H03 | i5MUDI-120 | TCCTGTTATG | CATAACAGGA
PlateB_A04 | i7MUDI-121 | GTACCTTGTT | PlateB_A04 | i5MUDI-121 | AAGTGCGATA | TATCGCACTT
PlateB_B04 | i7MUDI-122 | GAGTGCCTCT | PlateB_B04 | i5MUDI-122 | AGGCTACACG | CGTGTAGCCT
PlateB_C04 | i7MUDI-123 | TAAGTAGCGG | PlateB_C04 | i5MUDI-123 | CTATATCGGC | GCCGATATAG
PlateB_D04 | i7MUDI-124 | CGTGGTGTTC | PlateB_D04 | i5MUDI-124 | GCTAAGGTAA | TTACCTTAGC
PlateB_E04 | i7MUDI-125 | CATTCCTGAA | PlateB_E04 | i5MUDI-125 | TAACCTGGTT | AACCAGGTTA
PlateB_F04 | i7MUDI-126 | AAGATGCATG | PlateB_F04 | i5MUDI-126 | AGTTGGTCTA | TAGACCAACT
PlateB_G04 | i7MUDI-127 | CCTTGGAGCT | PlateB_G04 | i5MUDI-127 | ATGCAGCTGG | CCAGCTGCAT
PlateB_H04 | i7MUDI-128 | ACCGGAACAG | PlateB_H04 | i5MUDI-128 | CGTTGCCTTC | GAAGGCAACG
PlateB_A05 | i7MUDI-129 | GAATGGAAGC | PlateB_A05 | i5MUDI-129 | GCGTGGAGAA | TTCTCCACGC
PlateB_B05 | i7MUDI-130 | GTTCTCCATA | PlateB_B05 | i5MUDI-130 | TACGCCTCCT | AGGAGGCGTA
PlateB_C05 | i7MUDI-131 | GTCACTATGT | PlateB_C05 | i5MUDI-131 | AATTCGGTAG | CTACCGAATT
PlateB_D05 | i7MUDI-132 | TGGTAGAACT | PlateB_D05 | i5MUDI-132 | ATTGTCGAAC | GTTCGACAAT
PlateB_E05 | i7MUDI-133 | ACGCCTATGG | PlateB_E05 | i5MUDI-133 | CAACCTTGCG | CGCAAGGTTG
PlateB_F05 | i7MUDI-134 | AATCCGTTAC | PlateB_F05 | i5MUDI-134 | GCACTGCGTA | TACGCAGTGC
PlateB_G05 | i7MUDI-135 | GTTGAGGCTA | PlateB_G05 | i5MUDI-135 | TGCTAGTAGT | ACTACTAGCA
PlateB_H05 | i7MUDI-136 | TATCAACTGG | PlateB_H05 | i5MUDI-136 | AAGTCACGGA | TCCGTGACTT
PlateB_A06 | i7MUDI-137 | AAGAGGAGAT | PlateB_A06 | i5MUDI-137 | AGCGATTGAA | TTCAATCGCT
PlateB_B06 | i7MUDI-138 | GTCTTCTCGG | PlateB_B06 | i5MUDI-138 | CTACCTCTCT | AGAGAGGTAG
PlateB_C06 | i7MUDI-139 | GAAGCCACTC | PlateB_C06 | i5MUDI-139 | GACAACTGTC | GACAGTTGTC
PlateB_D06 | i7MUDI-140 | GTAGGACACA | PlateB_D06 | i5MUDI-140 | TCCATTGCGG | CCGCAATGGA
PlateB_E06 | i7MUDI-141 | CTCCTCGTAT | PlateB_E06 | i5MUDI-141 | AGCCTCGCAA | TTGCGAGGCT
PlateB_F06 | i7MUDI-142 | CCACATGATT | PlateB_F06 | i5MUDI-142 | AATACAGGCT | AGCCTGTATT
PlateB_G06 | i7MUDI-143 | AGACGGTTGG | PlateB_G06 | i5MUDI-143 | CGGACCGTTA | TAACGGTCCG
PlateB_H06 | i7MUDI-144 | CTAGGTTGAC | PlateB_H06 | i5MUDI-144 | GCGCTTATGC | GCATAAGCGC
PlateB_A07 | i7MUDI-145 | AAGCGTACCA | PlateB_A07 | i5MUDI-145 | TTAACACGAG | CTCGTGTTAA
PlateB_B07 | i7MUDI-146 | TCATGTTGGT | PlateB_B07 | i5MUDI-146 | CGCCTCTAGA | TCTAGAGGCG
PlateB_C07 | i7MUDI-147 | TTGGAATGGT | PlateB_C07 | i5MUDI-147 | AATCGACCTT | AAGGTCGATT
PlateB_D07 | i7MUDI-148 | GTGTATGTTG | PlateB_D07 | i5MUDI-148 | CCGCAATAAC | GTTATTGCGG
PlateB_E07 | i7MUDI-149 | TCCTGTCAAC | PlateB_E07 | i5MUDI-149 | GTTCCAACGA | TCGTTGGAAC
PlateB_F07 | i7MUDI-150 | TAATCAGGCA | PlateB_F07 | i5MUDI-150 | TGTTAGACCG | CGGTCTAACA
PlateB_G07 | i7MUDI-151 | GTAGTGGATT | PlateB_G07 | i5MUDI-151 | AACCTCATAG | CTATGAGGTT
PlateB_H07 | i7MUDI-152 | AATTGCGCAT | PlateB_H07 | i5MUDI-152 | ATGAATCCAC | GTGGATTCAT
PlateB_A08 | i7MUDI-153 | GACAATAACG | PlateB_A08 | i5MUDI-153 | CGGCTTAATT | AATTAAGCCG
PlateB_B08 | i7MUDI-154 | ACAGTTAAGC | PlateB_B08 | i5MUDI-154 | GAGTTGCAGG | CCTGCAACTC
PlateB_C08 | i7MUDI-155 | AGCCACACTA | PlateB_C08 | i5MUDI-155 | TCCACGAACA | TGTTCGTGGA
PlateB_D08 | i7MUDI-156 | CAATCGTCTT | PlateB_D08 | i5MUDI-156 | TGACGGAGGA | TCCTCCGTCA
PlateB_E08 | i7MUDI-157 | AGGAGCTTGT | PlateB_E08 | i5MUDI-157 | AATGAGTACG | CGTACTCATT
PlateB_F08 | i7MUDI-158 | TTGAGCGGAG | PlateB_F08 | i5MUDI-158 | CGTCTTCCGA | TCGGAAGACG
PlateB_G08 | i7MUDI-159 | AGTAGCTCTC | PlateB_G08 | i5MUDI-159 | GACAGAGATT | AATCTCTGTC
PlateB_H08 | i7MUDI-160 | CACGCTGTCA | PlateB_H08 | i5MUDI-160 | TTACGCTAAC | GTTAGCGTAA
PlateB_A09 | i7MUDI-161 | AAGACCTCTT | PlateB_A09 | i5MUDI-161 | CTCCTCGAAG | CTTCGAGGAG
PlateB_B09 | i7MUDI-162 | GACCTCTTCT | PlateB_B09 | i5MUDI-162 | ATACCGCAGA | TCTGCGGTAT
PlateB_C09 | i7MUDI-163 | TACTTCCTTG | PlateB_C09 | i5MUDI-163 | CCTATCTGAT | ATCAGATAGG
PlateB_D09 | i7MUDI-164 | TGCGATACGC | PlateB_D09 | i5MUDI-164 | GATCGGTTAC | GTAACCGATC
PlateB_E09 | i7MUDI-165 | GCAGGCTTAA | PlateB_E09 | i5MUDI-165 | TGGTGAGGTG | CACCTCACCA
PlateB_F09 | i7MUDI-166 | TAAGCTTGTG | PlateB_F09 | i5MUDI-166 | AACCGGCGTA | TACGCCGGTT
PlateB_G09 | i7MUDI-167 | ATGGTCCGCT | PlateB_G09 | i5MUDI-167 | AATACCGATC | GATCGGTATT
PlateB_H09 | i7MUDI-168 | ATGTCAGAAG | PlateB_H09 | i5MUDI-168 | CGATACTCAA | TTGAGTATCG
PlateB_A10 | i7MUDI-169 | GACGAAGGTC | PlateB_A10 | i5MUDI-169 | GTAAGGCGGT | ACCGCCTTAC
PlateB_B10 | i7MUDI-170 | ATCACCGTGA | PlateB_B10 | i5MUDI-170 | TTCAAGGTCG | CGACCTTGAA
PlateB_C10 | i7MUDI-171 | GCTACAGTGT | PlateB_C10 | i5MUDI-171 | TATCCGAGTA | TACTCGGATA
PlateB_D10 | i7MUDI-172 | CGTCGAATAT | PlateB_D10 | i5MUDI-172 | AGCGCGCTTA | TAAGCGCGCT
PlateB_E10 | i7MUDI-173 | CAACCATCGG | PlateB_E10 | i5MUDI-173 | CCGGAGACAT | ATGTCTCCGG
PlateB_F10 | i7MUDI-174 | CGGTCCATTC | PlateB_F10 | i5MUDI-174 | GAGATAACTG | CAGTTATCTC
PlateB_G10 | i7MUDI-175 | AGAAGAGCCA | PlateB_G10 | i5MUDI-175 | TTGTAAGCGC | GCGCTTACAA
PlateB_H10 | i7MUDI-176 | CTATGCAATG | PlateB_H10 | i5MUDI-176 | CAAGAGGAGG | CCTCCTCTTG
PlateB_A11 | i7MUDI-177 | CACTGAACCG | PlateB_A11 | i5MUDI-177 | AACCTTAGGA | TCCTAAGGTT
PlateB_B11 | i7MUDI-178 | TACTGTGTGA | PlateB_B11 | i5MUDI-178 | CTGGCAACTC | GAGTTGCCAG
PlateB_C11 | i7MUDI-179 | GCATTCTGTT | PlateB_C11 | i5MUDI-179 | GAACTTGTTG | CAACAAGTTC
PlateB_D11 | i7MUDI-180 | CTCCGCTAAG | PlateB_D11 | i5MUDI-180 | TGTGCAAGAT | ATCTTGCACA
PlateB_E11 | i7MUDI-181 | TCGCTTGAGA | PlateB_E11 | i5MUDI-181 | AATCGAGAGA | TCTCTCGATT
PlateB_F11 | i7MUDI-182 | AACTAGCCTT | PlateB_F11 | i5MUDI-182 | AGCGTGTCAG | CTGACACGCT
PlateB_G11 | i7MUDI-183 | TTCGCTCAGG | PlateB_G11 | i5MUDI-183 | CTTGGTGATT | AATCACCAAG
PlateB_H11 | i7MUDI-184 | CTCTACAACA | PlateB_H11 | i5MUDI-184 | GAAGCAGCAA | TTGCTGCTTC
PlateB_A12 | i7MUDI-185 | TGAGTGTGTT | PlateB_A12 | i5MUDI-185 | TTCCGTCGAC | GTCGACGGAA
PlateB_B12 | i7MUDI-186 | TAGTTAGTCG | PlateB_B12 | i5MUDI-186 | CGAGATGCCA | TGGCATCTCG
PlateB_C12 | i7MUDI-187 | GCCTGATCCT | PlateB_C12 | i5MUDI-187 | AAGTTCGTGC | GCACGAACTT
PlateB_D12 | i7MUDI-188 | CGAGTACAGG | PlateB_D12 | i5MUDI-188 | CGTCCATAAG | CTTATGGACG
PlateB_E12 | i7MUDI-189 | GCCTAGATTA | PlateB_E12 | i5MUDI-189 | TTGTGGCATA | TATGCCACAA
PlateB_F12 | i7MUDI-770 | CCAATGGCAC | PlateB_F12 | i5MUDI-770 | TTCTTGGCAG | CTGCCAAGAA
PlateB_G12 | i7MUDI-191 | CCGTGCAAGA | PlateB_G12 | i5MUDI-191 | CATTCTACTG | CAGTAGAATG
PlateB_H12 | i7MUDI-192 | CTGGCTGGTT | PlateB_H12 | i5MUDI-192 | ATCGCCGTAG | CTACGGCGAT
PlateC_A01 | i7MUDI-193 | CGTTAGGATT | PlateC_A01 | i5MUDI-193 | ATCCTTACAC | GTGTAAGGAT
PlateC_B01 | i7MUDI-194 | TTCCATTACG | PlateC_B01 | i5MUDI-194 | CGCAAGGACT | AGTCCTTGCG
PlateC_C01 | i7MUDI-195 | TAGCGGTAAC | PlateC_C01 | i5MUDI-195 | GCTGGCGTTA | TAACGCCAGC
PlateC_D01 | i7MUDI-196 | GTAGCCAGGA | PlateC_D01 | i5MUDI-196 | TACTTAGAGG | CCTCTAAGTA
PlateC_E01 | i7MUDI-197 | AGGATACTCT | PlateC_E01 | i5MUDI-197 | ATGGCGATGC | GCATCGCCAT
PlateC_F01 | i7MUDI-198 | TATCCTCCAG | PlateC_F01 | i5MUDI-198 | CATTGGTGCG | CGCACCAATG
PlateC_G01 | i7MUDI-199 | TAAGTCGTTC | PlateC_G01 | i5MUDI-199 | GCGAGATATA | TATATCTCGC
PlateC_H01 | i7MUDI-200 | TCCGGATTGA | PlateC_H01 | i5MUDI-200 | TGACTGCTAT | ATAGCAGTCA
PlateC_A02 | i7MUDI-201 | ACGTCTTGTT | PlateC_A02 | i5MUDI-201 | AACGTCCGCT | AGCGGACGTT
PlateC_B02 | i7MUDI-202 | ATGAAGTGCG | PlateC_B02 | i5MUDI-202 | CGCACATGTC | GACATGTGCG
PlateC_C02 | i7MUDI-203 | CGATCACTGC | PlateC_C02 | i5MUDI-203 | GCACACCTGA | TCAGGTGTGC
PlateC_D02 | i7MUDI-204 | CCTATCGGAA | PlateC_D02 | i5MUDI-204 | TTGTCCAGAG | CTCTGGACAA
PlateC_E02 | i7MUDI-205 | CAGAGAGCTT | PlateC_E02 | i5MUDI-205 | AGCCTTCCTG | CAGGAAGGCT
PlateC_F02 | i7MUDI-206 | GCAACTTGCG | PlateC_F02 | i5MUDI-206 | CCTTACGCCA | TGGCGTAAGG
PlateC_G02 | i7MUDI-207 | TATGGAGGAC | PlateC_G02 | i5MUDI-207 | GAATACGTAC | GTACGTATTC
PlateC_H02 | i7MUDI-208 | TGAGATCAGA | PlateC_H02 | i5MUDI-208 | TTGGCACCGT | ACGGTGCCAA
PlateC_A03 | i7MUDI-209 | TCAGCCTATT | PlateC_A03 | i5MUDI-209 | ATTAGGTGGC | GCCACCTAAT
PlateC_B03 | i7MUDI-210 | GTTGTGAGCG | PlateC_B03 | i5MUDI-210 | CGATCAAGAA | TTCTTGATCG
PlateC_C03 | i7MUDI-211 | TCAGTAACAC | PlateC_C03 | i5MUDI-211 | GCTGTCTTCT | AGAAGACAGC
PlateC_D03 | i7MUDI-212 | AAGGCTCAGA | PlateC_D03 | i5MUDI-212 | TACATGTCTG | CAGACATGTA
PlateC_E03 | i7MUDI-213 | GTGTGGTGGT | PlateC_E03 | i5MUDI-213 | AACCAGTTGA | TCAACTGGTT
PlateC_F03 | i7MUDI-214 | CCGAGCTTAG | PlateC_F03 | i5MUDI-214 | CCGGTAAGCT | AGCTTACCGG
PlateC_G03 | i7MUDI-215 | ATCACGCTTC | PlateC_G03 | i5MUDI-215 | GTTCGAATAG | CTATTCGAAC
PlateC_H03 | i7MUDI-216 | TAGCTATGCA | PlateC_H03 | i5MUDI-216 | TGTCAGGCTC | GAGCCTGACA
PlateC_A04 | i7MUDI-217 | TGTTCCTCAT | PlateC_A04 | i5MUDI-217 | CAACAGTGTT | AACACTGTTG
PlateC_B04 | i7MUDI-218 | CATACCTTCT | PlateC_B04 | i5MUDI-218 | AAGAGAGGAA | TTCCTCTCTT
PlateC_C04 | i7MUDI-219 | GCCTTCAATG | PlateC_C04 | i5MUDI-219 | CGGTTGTAGC | GCTACAACCG
PlateC_D04 | i7MUDI-220 | CTTGACCAGC | PlateC_D04 | i5MUDI-220 | GCCTGAAGTG | CACTTCAGGC
PlateC_E04 | i7MUDI-221 | CTACACACAA | PlateC_E04 | i5MUDI-221 | TTACGACACT | AGTGTCGTAA
PlateC_F04 | i7MUDI-222 | TAGGCTGAAT | PlateC_F04 | i5MUDI-222 | CGCCTAGATC | GATCTAGGCG
PlateC_G04 | i7MUDI-223 | TCGGAGTCCT | PlateC_G04 | i5MUDI-223 | AATCTGGATG | CATCCAGATT
PlateC_H04 | i7MUDI-224 | AACATCGCGG | PlateC_H04 | i5MUDI-224 | CGACGGTACA | TGTACCGTCG
PlateC_A05 | i7MUDI-225 | GTTGTCTTAC | PlateC_A05 | i5MUDI-225 | GTAGTATTGC | GCAATACTAC
PlateC_B05 | i7MUDI-226 | GTGGCAACTA | PlateC_B05 | i5MUDI-226 | TCCAGCGGAT | ATCCGCTGGA
PlateC_C05 | i7MUDI-227 | GAGCAGGCAT | PlateC_C05 | i5MUDI-227 | CAACCACCTC | GAGGTGGTTG
PlateC_D05 | i7MUDI-228 | AACGGCACCT | PlateC_D05 | i5MUDI-228 | AGCTTAGGCG | CGCCTAAGCT
PlateC_E05 | i7MUDI-229 | AGTAACCTTG | PlateC_E05 | i5MUDI-229 | CCGGTTCCTT | AAGGAACCGG
PlateC_F05 | i7MUDI-230 | TCTCATAAGC | PlateC_F05 | i5MUDI-230 | GACATTGAAC | GTTCAATGTC
PlateC_G05 | i7MUDI-231 | TGCTTGCCAA | PlateC_G05 | i5MUDI-231 | TTAGAGGCGA | TCGCCTCTAA
PlateC_H05 | i7MUDI-232 | CGGTTCCTGT | PlateC_H05 | i5MUDI-232 | CAAGCCGAAC | GTTCGGCTTG
PlateC_A06 | i7MUDI-233 | CCAAGTAGAT | PlateC_A06 | i5MUDI-233 | AGGAGAACGG | CCGTTCTCCT
PlateC_B06 | i7MUDI-234 | AAGGTTGGCG | PlateC_B06 | i5MUDI-234 | CCTGTTAGAC | GTCTAACAGG
PlateC_C06 | i7MUDI-235 | TGCTCTGGTC | PlateC_C06 | i5MUDI-235 | GTTCTACGTT | AACGTAGAAC
PlateC_D06 | i7MUDI-236 | ACTGTAACGA | PlateC_D06 | i5MUDI-236 | TAAGTCCACA | TGTGGACTTA
PlateC_E06 | i7MUDI-237 | GATTCCAGGT | PlateC_E06 | i5MUDI-237 | CAAGAACCAT | ATGGTTCTTG
PlateC_F06 | i7MUDI-238 | TTCACCAGAT | PlateC_F06 | i5MUDI-238 | AGTTGATGAC | GTCATCAACT
PlateC_G06 | i7MUDI-239 | ACTTCCAAGG | PlateC_G06 | i5MUDI-239 | CCTACTCTTG | CAAGAGTAGG
PlateC_H06 | i7MUDI-240 | CCGAATATTC | PlateC_H06 | i5MUDI-240 | GAACAATCCA | TGGATTGTTC
PlateC_A07 | i7MUDI-241 | CTCTTATCCA | PlateC_A07 | i5MUDI-241 | TTCTGTTGGT | ACCAACAGAA
PlateC_B07 | i7MUDI-242 | TCACAGCGGT | PlateC_B07 | i5MUDI-242 | CATCGTCAGG | CCTGACGATG
PlateC_C07 | i7MUDI-243 | CCTCTGTCGT | PlateC_C07 | i5MUDI-243 | ATGCATGAAG | CTTCATGCAT
PlateC_D07 | i7MUDI-244 | TCTGTTCTCG | PlateC_D07 | i5MUDI-244 | CGTGAATCGC | GCGATTCACG
PlateC_E07 | i7MUDI-245 | GATACTTCAC | PlateC_E07 | i5MUDI-245 | GAGCAGCCTT | AAGGCTGCTC
PlateC_F07 | i7MUDI-246 | AGTGCTGATA | PlateC_F07 | i5MUDI-246 | TCGATTACCA | TGGTAATCGA
PlateC_G07 | i7MUDI-247 | ATCCTTCGGT | PlateC_G07 | i5MUDI-247 | CAGTCCAATT | AATTGGACTG
PlateC_H07 | i7MUDI-248 | GACAACGATT | PlateC_H07 | i5MUDI-248 | AGAGGCTTGG | CCAAGCCTCT
PlateC_A08 | i7MUDI-249 | GAACCGGTAG | PlateC_A08 | i5MUDI-249 | CAGGCTCTCA | TGAGAGCCTG
PlateC_B08 | i7MUDI-250 | AGCAATGAGC | PlateC_B08 | i5MUDI-250 | GTTCGCTCTC | GAGAGCGAAC
PlateC_C08 | i7MUDI-251 | CAAGACTCCA | PlateC_C08 | i5MUDI-251 | TCGGACTAAT | ATTAGTCCGA
PlateC_D08 | i7MUDI-252 | ACCGTGTAGG | PlateC_D08 | i5MUDI-252 | CGAGATCTTC | GAAGATCTCG
PlateC_E08 | i7MUDI-253 | AGGCACAGGT | PlateC_E08 | i5MUDI-253 | ATAACCGGAC | GTCCGGTTAT
PlateC_F08 | i7MUDI-254 | CGACAGATCG | PlateC_F08 | i5MUDI-254 | CGTGTAGTTA | TAACTACACG
PlateC_G08 | i7MUDI-255 | ACGCGACAAC | PlateC_G08 | i5MUDI-255 | GAACATAGGT | ACCTATGTTC
PlateC_H08 | i7MUDI-256 | ACTTGCGTTA | PlateC_H08 | i5MUDI-256 | TCTAACATCG | CGATGTTAGA
PlateC_A09 | i7MUDI-257 | CACCACTCAT | PlateC_A09 | i5MUDI-257 | AACGGTGGCA | TGCCACCGTT
PlateC_B09 | i7MUDI-258 | CTTCGTAACT | PlateC_B09 | i5MUDI-258 | AGGACGGTGT | ACACCGTCCT
PlateC_C09 | i7MUDI-259 | CAGTATTCGG | PlateC_C09 | i5MUDI-259 | CTGTGACCTG | CAGGTCACAG
PlateC_D09 | i7MUDI-260 | CAGTCTGGAC | PlateC_D09 | i5MUDI-260 | GCTGTAACAA | TTGTTACAGC
PlateC_E09 | i7MUDI-261 | TACCGTTCTA | PlateC_E09 | i5MUDI-261 | TACGGACGTC | GACGTCCGTA
PlateC_F09 | i7MUDI-262 | GTGTCCACAG | PlateC_F09 | i5MUDI-262 | CCTAAGGAGC | GCTCCTTAGG
PlateC_G09 | i7MUDI-263 | TTACGACTGT | PlateC_G09 | i5MUDI-263 | ATAAGGCCAG | CTGGCCTTAT
PlateC_H09 | i7MUDI-264 | GACGCGAATG | PlateC_H09 | i5MUDI-264 | CTCATCTGTA | TACAGATGAG
PlateC_A10 | i7MUDI-265 | CAACGTACGC | PlateC_A10 | i5MUDI-265 | GAAGGCATCT | AGATGCCTTC
PlateC_B10 | i7MUDI-266 | AGCTCAGGAA | PlateC_B10 | i5MUDI-266 | TCTCTACTGC | GCAGTAGAGA
PlateC_C10 | i7MUDI-267 | GATAGGCGGT | PlateC_C10 | i5MUDI-267 | AACCGAACAA | TTGTTCGGTT
PlateC_D10 | i7MUDI-268 | AGTAGGAAGT | PlateC_D10 | i5MUDI-268 | ATCTCGCCAC | GTGGCGAGAT
PlateC_E10 | i7MUDI-269 | CATGTTGTAG | PlateC_E10 | i5MUDI-269 | CCATGCAACG | CGTTGCATGG
PlateC_F10 | i7MUDI-270 | CACATTCTTC | PlateC_F10 | i5MUDI-270 | GAATGGTGTA | TACACCATTC
PlateC_G10 | i7MUDI-271 | GCAGCTCGTA | PlateC_G10 | i5MUDI-271 | TATATGCCGT | ACGGCATATA
PlateC_H10 | i7MUDI-272 | GTTCAGACGG | PlateC_H10 | i5MUDI-272 | CTCGATAGAT | ATCTATCGAG
PlateC_A11 | i7MUDI-273 | TCCTGGAAGT | PlateC_A11 | i5MUDI-273 | AACACAAGAG | CTCTTGTGTT
PlateC_B11 | i7MUDI-274 | GCATTGTTAG | PlateC_B11 | i5MUDI-274 | CGCAATCGGT | ACCGATTGCG
PlateC_C11 | i7MUDI-275 | GACCTACAGC | PlateC_C11 | i5MUDI-275 | GTTGCGTAGA | TCTACGCAAC
PlateC_D11 | i7MUDI-276 | CACCGACGTA | PlateC_D11 | i5MUDI-276 | TAGAGTGATC | GATCACTCTA
PlateC_E11 | i7MUDI-277 | CTCTCACCTT | PlateC_E11 | i5MUDI-277 | AAGACGCAGC | GCTGCGTCTT
PlateC_F11 | i7MUDI-278 | CTCGTTCATT | PlateC_F11 | i5MUDI-278 | AACTTCTCGA | TCGAGAAGTT
PlateC_G11 | i7MUDI-279 | TGGTGGCAAG | PlateC_G11 | i5MUDI-279 | CGCAACTGAG | CTCAGTTGCG
PlateC_H11 | i7MUDI-280 | GATTGCTTGA | PlateC_H11 | i5MUDI-280 | GCTCCGCAAT | ATTGCGGAGC
PlateC_A12 | i7MUDI-281 | CCGTTAAGGT | PlateC_A12 | i5MUDI-281 | GTAACTTCCG | CGGAAGTTAC
PlateC_B12 | i7MUDI-282 | TGCTGAGAGG | PlateC_B12 | i5MUDI-282 | CTCACGACTA | TAGTCGTGAG
PlateC_C12 | i7MUDI-283 | TTGTCACTTG | PlateC_C12 | i5MUDI-283 | AACCAACGGC | GCCGTTGGTT
PlateC_D12 | i7MUDI-284 | GCTGTTATGT | PlateC_D12 | i5MUDI-284 | CCTGCCTGTA | TACAGGCAGG
PlateC_E12 | i7MUDI-285 | GCAGCAGTTG | PlateC_E12 | i5MUDI-285 | TACGCTGCAG | CTGCAGCGTA
PlateC_F12 | i7MUDI-286 | GCAGATCAAT | PlateC_F12 | i5MUDI-286 | AATGTTGCGA | TCGCAACATT
PlateC_G12 | i7MUDI-287 | TGGTTCACGG | PlateC_G12 | i5MUDI-287 | CGACGTTCTG | CAGAACGTCG
PlateC_H12 | i7MUDI-288 | TCGACCGCAT | PlateC_H12 | i5MUDI-288 | AATAGGACAC | GTGTCCTATT
PlateD_A01 | i7MUDI-289 | TAACCTAGGT | PlateD_A01 | i5MUDI-289 | ATGTGCCTCA | TGAGGCACAT
PlateD_B01 | i7MUDI-290 | AACTCATGCG | PlateD_B01 | i5MUDI-290 | CGACTCCGTT | AACGGAGTCG
PlateD_C01 | i7MUDI-291 | CCGGATGAAC | PlateD_C01 | i5MUDI-291 | GCTGTTGTGG | CCACAACAGC
PlateD_D01 | i7MUDI-292 | CGTTGCCGTA | PlateD_D01 | i5MUDI-292 | TACCAATCAC | GTGATTGGTA
PlateD_E01 | i7MUDI-293 | GCTCTACGGT | PlateD_E01 | i5MUDI-293 | ATGTCTTACG | CGTAAGACAT
PlateD_F01 | i7MUDI-294 | TGCATTGGCG | PlateD_F01 | i5MUDI-294 | CGCAACAATA | TATTGTTGCG
PlateD_G01 | i7MUDI-295 | CGATTGTGAC | PlateD_G01 | i5MUDI-295 | GAACGAAGAC | GTCTTCGTTC
PlateD_H01 | i7MUDI-296 | GACTGCACTA | PlateD_H01 | i5MUDI-296 | TCGAGGACGT | ACGTCCTCGA
PlateD_A02 | i7MUDI-297 | GTTAACTGCT | PlateD_A02 | i5MUDI-297 | ATTATGAGCG | CGCTCATAAT
PlateD_B02 | i7MUDI-298 | TCGGACCTTG | PlateD_B02 | i5MUDI-298 | CGCGTTATAA | TTATAACGCG
PlateD_C02 | i7MUDI-299 | TGCAGCAAGC | PlateD_C02 | i5MUDI-299 | GCGTGCATGT | ACATGCACGC
PlateD_D02 | i7MUDI-300 | CACATGCGAA | PlateD_D02 | i5MUDI-300 | TAAGCGGCTC | GAGCCGCTTA
PlateD_E02 | i7MUDI-301 | CAGACGTAAT | PlateD_E02 | i5MUDI-301 | AACATGGAGA | TCTCCATGTT
PlateD_F02 | i7MUDI-302 | ATTCGGTACG | PlateD_F02 | i5MUDI-302 | CCGAGTCTCT | AGAGACTCGG
PlateD_G02 | i7MUDI-303 | TTAGCACGGC | PlateD_G02 | i5MUDI-303 | GTACTTCTAC | GTAGAAGTAC
PlateD_H02 | i7MUDI-304 | GAGGATAGTA | PlateD_H02 | i5MUDI-304 | TGTTCACATG | CATGTGAACA
PlateD_A03 | i7MUDI-305 | AACTGTGGTT | PlateD_A03 | i5MUDI-305 | AAGGTAACGC | GCGTTACCTT
PlateD_B03 | i7MUDI-306 | ATTACCTCGG | PlateD_B03 | i5MUDI-306 | CCGCCTTACT | AGTAAGGCGG
PlateD_C03 | i7MUDI-307 | CGCTGTATAC | PlateD_C03 | i5MUDI-307 | GTTGAGGCAG | CTGCCTCAAC
PlateD_D03 | i7MUDI-308 | CTTGCTCACA | PlateD_D03 | i5MUDI-308 | TGGCGACCTA | TAGGTCGCCA
PlateD_E03 | i7MUDI-309 | CAACACCTGT | PlateD_E03 | i5MUDI-309 | AGAAGCGACA | TGTCGCTTCT
PlateD_F03 | i7MUDI-310 | CAATTGCTCG | PlateD_F03 | i5MUDI-310 | CAGGATAATC | GATTATCCTG
PlateD_G03 | i7MUDI-311 | CATAGACAAC | PlateD_G03 | i5MUDI-311 | GCTCCTACAG | CTGTAGGAGC
PlateD_H03 | i7MUDI-312 | TTGGTGTCTA | PlateD_H03 | i5MUDI-312 | TTCAACAGGT | ACCTGTTGAA
PlateD_A04 | i7MUDI-313 | TATGTCCTGT | PlateD_A04 | i5MUDI-313 | CCTCGTCCAT | ATGGACGAGG
PlateD_B04 | i7MUDI-314 | GCCAATTCGT | PlateD_B04 | i5MUDI-314 | AGCGTTGGTT | AACCAACGCT
PlateD_C04 | i7MUDI-315 | TAGGCGATCG | PlateD_C04 | i5MUDI-315 | CATTCGAACA | TGTTCGAATG
PlateD_D04 | i7MUDI-316 | ATGAGTGTAC | PlateD_D04 | i5MUDI-316 | GCTTACCGAC | GTCGGTAAGC
PlateD_E04 | i7MUDI-317 | CCGAAGGATA | PlateD_E04 | i5MUDI-317 | TTAGCTTAGG | CCTAAGCTAA
PlateD_F04 | i7MUDI-318 | AGTCCACTGT | PlateD_F04 | i5MUDI-318 | CCGACACACA | TGTGTGTCGG
PlateD_G04 | i7MUDI-319 | GCGGCTAATT | PlateD_G04 | i5MUDI-319 | ATTCGCTGAT | ATCAGCGAAT
PlateD_H04 | i7MUDI-320 | TCTAACTCAG | PlateD_H04 | i5MUDI-320 | CCAAGAGGCA | TGCCTCTTGG
PlateD_A05 | i7MUDI-321 | CAAGCTGAGC | PlateD_A05 | i5MUDI-321 | GACGCAGTTC | GAACTGCGTC
PlateD_B05 | i7MUDI-322 | CCAGAGCACA | PlateD_B05 | i5MUDI-322 | TGGAACTCGG | CCGAGTTCCA
PlateD_C05 | i7MUDI-323 | TGTACAAGGT | PlateD_C05 | i5MUDI-323 | CCACACCAAT | ATTGGTGTGG
PlateD_D05 | i7MUDI-324 | TAGAATGCCT | PlateD_D05 | i5MUDI-324 | AGTTCTCGGC | GCCGAGAACT
PlateD_E05 | i7MUDI-325 | TGTCTTACTG | PlateD_E05 | i5MUDI-325 | CTTGACGACG | CGTCGTCAAG
PlateD_F05 | i7MUDI-326 | ATGACTAAGC | PlateD_F05 | i5MUDI-326 | GAGGTCGCTA | TAGCGACCTC
PlateD_G05 | i7MUDI-327 | ATGTAGGCAA | PlateD_G05 | i5MUDI-327 | TCAGTAGCAT | ATGCTACTGA
PlateD_H05 | i7MUDI-328 | GCGAAGAGGT | PlateD_H05 | i5MUDI-328 | CTAACGTGGA | TCCACGTTAG
PlateD_A06 | i7MUDI-329 | CGGTGGTTCT | PlateD_A06 | i5MUDI-329 | ATGCCAACCG | CGGTTGGCAT
PlateD_B06 | i7MUDI-330 | CTGTCGTTGG | PlateD_B06 | i5MUDI-330 | CGGTCGATTC | GAATCGACCG
PlateD_C06 | i7MUDI-331 | TGATCGACAC | PlateD_C06 | i5MUDI-331 | GAAGTACAGT | ACTGTACTTC
PlateD_D06 | i7MUDI-332 | CCACCAGCTA | PlateD_D06 | i5MUDI-332 | TCTGCAGTAA | TTACTGCAGA
PlateD_E06 | i7MUDI-333 | CACGGTTCGT | PlateD_E06 | i5MUDI-333 | CTATCCTAGC | GCTAGGATAG
PlateD_F06 | i7MUDI-334 | AGTGAGAGCT | PlateD_F06 | i5MUDI-334 | AACACTCCTT | AAGGAGTGTT
PlateD_G06 | i7MUDI-335 | TTGCATGCGG | PlateD_G06 | i5MUDI-335 | CCGAACCTAA | TTAGGTTCGG
PlateD_H06 | i7MUDI-336 | TATACGTGTC | PlateD_H06 | i5MUDI-336 | GTCTAGTCGC | GCGACTAGAC
PlateD_A07 | i7MUDI-337 | TGACGCGTTA | PlateD_A07 | i5MUDI-337 | TGGATGTACG | CGTACATCCA
PlateD_B07 | i7MUDI-338 | TACAGAACGT | PlateD_B07 | i5MUDI-338 | CTACCAGCGT | ACGCTGGTAG
PlateD_C07 | i7MUDI-339 | CTTGTCAGGT | PlateD_C07 | i5MUDI-339 | AAGGATTCAG | CTGAATCCTT
PlateD_D07 | i7MUDI-340 | ATCCACAGCG | PlateD_D07 | i5MUDI-340 | CGAGGTGTGT | ACACACCTCG
PlateD_E07 | i7MUDI-341 | CCTATCCATC | PlateD_E07 | i5MUDI-341 | GTAGACGCTC | GAGCGTCTAC
PlateD_F07 | i7MUDI-342 | ACCGCGAGTA | PlateD_F07 | i5MUDI-342 | TCGTCCGTCA | TGACGGACGA
PlateD_G07 | i7MUDI-343 | AAGTTCTGGT | PlateD_G07 | i5MUDI-343 | CCGTGATAGG | CCTATCACGG
PlateD_H07 | i7MUDI-344 | ACAGGTATCG | PlateD_H07 | i5MUDI-344 | AGGATGACCT | AGGTCATCCT
PlateD_A08 | i7MUDI-345 | ATGACGGATT | PlateD_A08 | i5MUDI-345 | CCTCGAGTAC | GTACTCGAGG
PlateD_B08 | i7MUDI-346 | GTCTGAGTAG | PlateD_B08 | i5MUDI-346 | GTCACTGAGG | CCTCAGTGAC
PlateD_C08 | i7MUDI-347 | TGCCAGATGT | PlateD_C08 | i5MUDI-347 | TACGGTTAGA | TCTAACCGTA
PlateD_D08 | i7MUDI-348 | GCTAAGCATT | PlateD_D08 | i5MUDI-348 | CAACGAGAAT | ATTCTCGTTG
PlateD_E08 | i7MUDI-349 | ACAGCATGGT | PlateD_E08 | i5MUDI-349 | AATACACCGG | CCGGTGTATT
PlateD_F08 | i7MUDI-350 | ATAGAGACCG | PlateD_F08 | i5MUDI-350 | CCGATCCATC | GATGGATCGG
PlateD_G08 | i7MUDI-351 | ATATCGCGTA | PlateD_G08 | i5MUDI-351 | GAATCTCGCT | AGCGAGATTC
PlateD_H08 | i7MUDI-352 | TTAAGGAGGT | PlateD_H08 | i5MUDI-352 | TGACCGGCAA | TTGCCGGTCA
PlateD_A09 | i7MUDI-353 | CTGTGCGACT | PlateD_A09 | i5MUDI-353 | CATGATAGCA | TGCTATCATG
PlateD_B09 | i7MUDI-354 | TCCGTATGCT | PlateD_B09 | i5MUDI-354 | AACAGCTTCG | CGAAGCTGTT
PlateD_C09 | i7MUDI-355 | CCATCGATGT | PlateD_C09 | i5MUDI-355 | CTAGTGCTTA | TAAGCACTAG
PlateD_D09 | i7MUDI-356 | GTGAGCCGTT | PlateD_D09 | i5MUDI-356 | TGTGATACGT | ACGTATCACA
PlateD_E09 | i7MUDI-357 | TGCCGTTAAT | PlateD_E09 | i5MUDI-357 | ATGAGCGTAT | ATACGCTCAT
PlateD_F09 | i7MUDI-358 | CGGATGTGGT | PlateD_F09 | i5MUDI-358 | CTAGATATGG | CCATATCTAG
PlateD_G09 | i7MUDI-359 | TCGCGTGTTG | PlateD_G09 | i5MUDI-359 | CGCTATGCTG | CAGCATAGCG
PlateD_H09 | i7MUDI-360 | CCGCGATCAT | PlateD_H09 | i5MUDI-360 | TACTACGTGA | TCACGTAGTA
PlateD_A10 | i7MUDI-361 | CGCGTTATCG | PlateD_A10 | i5MUDI-361 | ATGTGGAGGT | ACCTCCACAT
PlateD_B10 | i7MUDI-362 | GTAGCCTCCT | PlateD_B10 | i5MUDI-362 | CCATGGCTCA | TGAGCCATGG
PlateD_C10 | i7MUDI-363 | ACTAGACACT | PlateD_C10 | i5MUDI-363 | CCAATCACGC | GCGTGATTGG
PlateD_D10 | i7MUDI-364 | CGATTCGTTG | PlateD_D10 | i5MUDI-364 | TTAGATCCAG | CTGGATCTAA
PlateD_E10 | i7MUDI-365 | GAAGAGATGT | PlateD_E10 | i5MUDI-365 | AGGAATATCG | CGATATTCCT
PlateD_F10 | i7MUDI-366 | AGATCCGACG | PlateD_F10 | i5MUDI-366 | CCTCCTATGT | ACATAGGAGG
PlateD_G10 | i7MUDI-367 | CCAGGACATT | PlateD_G10 | i5MUDI-367 | TAGAGACACG | CGTGTCTCTA
PlateD_H10 | i7MUDI-368 | ACGTGGCATT | PlateD_H10 | i5MUDI-368 | CCAGCTCAGT | ACTGAGCTGG
PlateD_A11 | i7MUDI-369 | AAGCAGGACG | PlateD_A11 | i5MUDI-369 | ATGGCTCATA | TATGAGCCAT
PlateD_B11 | i7MUDI-370 | ACGAGTCGGT | PlateD_B11 | i5MUDI-370 | CGGAGTGAAG | CTTCACTCCG
PlateD_C11 | i7MUDI-371 | AGTGTACGCG | PlateD_C11 | i5MUDI-371 | TACCTATGGT | ACCATAGGTA
PlateD_D11 | i7MUDI-372 | ACCGACCATT | PlateD_D11 | i5MUDI-372 | ATGAGACAGT | ACTGTCTCAT
PlateD_E11 | i7MUDI-373 | TTGCTAACGT | PlateD_E11 | i5MUDI-373 | CTAAGAGTTG | CAACTCTTAG
PlateD_F11 | i7MUDI-374 | CTTGATACTG | PlateD_F11 | i5MUDI-374 | TAACCGTATG | CATACGGTTA
PlateD_G11 | i7MUDI-375 | CTGGATAAGT | PlateD_G11 | i5MUDI-375 | AGAGTCCATG | CATGGACTCT
PlateD_H11 | i7MUDI-771 | ACCGAAGCCA | PlateD_H11 | i5MUDI-771 | GACTCTTCCA | TGGAAGAGTC
PlateD_A12 | i7MUDI-377 | GTCCATGAGT | PlateD_A12 | i5MUDI-377 | TATGGCTTGT | ACAAGCCATA
PlateD_B12 | i7MUDI-378 | ACTCCAGTCG | PlateD_B12 | i5MUDI-378 | CGTTGTTCCT | AGGAACAACG
PlateD_C12 | i7MUDI-379 | TCTCAGCACG | PlateD_C12 | i5MUDI-379 | CCGACATTAG | CTAATGTCGG
PlateD_D12 | i7MUDI-380 | ATCGTGATGT | PlateD_D12 | i5MUDI-380 | TGTGAAGGCA | TGCCTTCACA
PlateD_E12 | i7MUDI-772 | TGGAAGAGTC | PlateD_E12 | i5MUDI-772 | TGGCTTCGGT | ACCGAAGCCA
PlateD_F12 | i7MUDI-382 | GAGATCGGCT | PlateD_F12 | i5MUDI-382 | CCGACTAGGA | TCCTAGTCGG
PlateD_G12 | i7MUDI-383 | CTACGTCTCG | PlateD_G12 | i5MUDI-383 | AACATTACCG | CGGTAATGTT
PlateD_H12 | i7MUDI-384 | CTCAGGCTGT | PlateD_H12 | i5MUDI-384 | CCTAATTCGT | ACGAATTAGG
PlateE_A01 | i7MUDI-385 | CAGCGTTGTT | PlateE_A01 | i5MUDI-385 | AAGGCACAAT | ATTGTGCCTT
PlateE_B01 | i7MUDI-386 | ACCTTAGACG | PlateE_B01 | i5MUDI-386 | CGGTTCCACA | TGTGGAACCG
PlateE_C01 | i7MUDI-387 | AGGTACCAGC | PlateE_C01 | i5MUDI-387 | GCCAATATTC | GAATATTGGC
PlateE_D01 | i7MUDI-388 | CAACCGCCAA | PlateE_D01 | i5MUDI-388 | TTAATGGAGG | CCTCCATTAA
PlateE_E01 | i7MUDI-389 | CAATACAGCT | PlateE_E01 | i5MUDI-389 | ATGTAGGCCT | AGGCCTACAT
PlateE_F01 | i7MUDI-390 | TCATGACCGG | PlateE_F01 | i5MUDI-390 | CATAGGTATC | GATACCTATG
PlateE_G01 | i7MUDI-391 | TCGACTTATC | PlateE_G01 | i5MUDI-391 | GCCATCAGAG | CTCTGATGGC
PlateE_H01 | i7MUDI-392 | GTTGGAAGAA | PlateE_H01 | i5MUDI-392 | TGGTCTTAGA | TCTAAGACCA
PlateE_A02 | i7MUDI-393 | CCAGAATTCT | PlateE_A02 | i5MUDI-393 | ATTGGACTTG | CAAGTCCAAT
PlateE_B02 | i7MUDI-394 | TTACCAACAG | PlateE_B02 | i5MUDI-394 | CCATTATCGA | TCGATAATGG
PlateE_C02 | i7MUDI-395 | TTGAAGCCGC | PlateE_C02 | i5MUDI-395 | GAACTGTACT | AGTACAGTTC
PlateE_D02 | i7MUDI-396 | AAGAACGGTA | PlateE_D02 | i5MUDI-396 | TGCAGGAGAC | GTCTCCTGCA
PlateE_E02 | i7MUDI-397 | TATCTGAGCT | PlateE_E02 | i5MUDI-397 | AACCGGAAGG | CCTTCCGGTT
PlateE_F02 | i7MUDI-398 | AAGCTGTCGG | PlateE_F02 | i5MUDI-398 | CTGGACCAAC | GTTGGTCCAG
PlateE_G02 | i7MUDI-399 | GTGTGATAAC | PlateE_G02 | i5MUDI-399 | GCTTCGAGCT | AGCTCGAAGC
PlateE_H02 | i7MUDI-400 | AACGGTCTTA | PlateE_H02 | i5MUDI-400 | TGAGGCTGTA | TACAGCCTCA
PlateE_A03 | i7MUDI-401 | AACCTTGCAT | PlateE_A03 | i5MUDI-401 | ACCGGTATAT | ATATACCGGT
PlateE_B03 | i7MUDI-402 | AATAGCCGCG | PlateE_B03 | i5MUDI-402 | CTCCTTGCCA | TGGCAAGGAG
PlateE_C03 | i7MUDI-403 | CCTTGGCTTC | PlateE_C03 | i5MUDI-403 | GAACTCAATC | GATTGAGTTC
PlateE_D03 | i7MUDI-404 | GTAGTTCTGA | PlateE_D03 | i5MUDI-404 | TGTTAAGTGG | CCACTTAACA
PlateE_E03 | i7MUDI-405 | GCTTGTGAAT | PlateE_E03 | i5MUDI-405 | AGGAACTGCT | AGCAGTTCCT
PlateE_F03 | i7MUDI-406 | AGAGTCAATG | PlateE_F03 | i5MUDI-406 | CCAATAGAAC | GTTCTATTGG
PlateE_G03 | i7MUDI-773 | CTGCCAAGAA | PlateE_G03 | i5MUDI-773 | GTGCCATTGG | CCAATGGCAC
PlateE_H03 | i7MUDI-408 | AACACGAACA | PlateE_H03 | i5MUDI-408 | TAACCACTGG | CCAGTGGTTA
PlateE_A04 | i7MUDI-409 | GCCTACTGAG | PlateE_A04 | i5MUDI-409 | ACCTACCATT | AATGGTAGGT
PlateE_B04 | i7MUDI-410 | CCTAACGCCT | PlateE_B04 | i5MUDI-410 | AGTGGTTACA | TGTAACCACT
PlateE_C04 | i7MUDI-411 | TCATCCGCTG | PlateE_C04 | i5MUDI-411 | CCTTGCGGTT | AACCGCAAGG
PlateE_D04 | i7MUDI-412 | AAGTTCCTAC | PlateE_D04 | i5MUDI-412 | GTTCCAGAAC | GTTCTGGAAC
PlateE_E04 | i7MUDI-413 | AGCATCTAGA | PlateE_E04 | i5MUDI-413 | TAGGTTCCGG | CCGGAACCTA
PlateE_F04 | i7MUDI-414 | CCTGGTTACG | PlateE_F04 | i5MUDI-414 | ACAGCCGCTT | AAGCGGCTGT
PlateE_G04 | i7MUDI-415 | CACGCAAGAT | PlateE_G04 | i5MUDI-415 | AACAACCGAA | TTCGGTTGTT
PlateE_H04 | i7MUDI-416 | GATAAGTTGG | PlateE_H04 | i5MUDI-416 | CCGTCTTGTG | CACAAGACGG
PlateE_A05 | i7MUDI-417 | GCCAACCTTC | PlateE_A05 | i5MUDI-417 | GTGTGTAAGC | GCTTACACAC
PlateE_B05 | i7MUDI-418 | CGACAAGGCA | PlateE_B05 | i5MUDI-418 | TGGTGAATCT | AGATTCACCA
PlateE_C05 | i7MUDI-419 | TAGTGCTAGG | PlateE_C05 | i5MUDI-419 | CAAGTCTGGT | ACCAGACTTG
PlateE_D05 | i7MUDI-420 | TGGAGCGATT | PlateE_D05 | i5MUDI-420 | AACTTGGCTC | GAGCCAAGTT
PlateE_E05 | i7MUDI-421 | CAGCGGATAG | PlateE_E05 | i5MUDI-421 | CGGTAGAGGA | TCCTCTACCG
PlateE_F05 | i7MUDI-422 | CTTCCTATGC | PlateE_F05 | i5MUDI-422 | GCTTAATACG | CGTATTAAGC
PlateE_G05 | i7MUDI-423 | TTGAACTCCA | PlateE_G05 | i5MUDI-423 | TTGGTGCGAT | ATCGCACCAA
PlateE_H05 | i7MUDI-424 | GTGAGGAATG | PlateE_H05 | i5MUDI-424 | ACGAAGGCGA | TCGCCTTCGT
PlateE_A06 | i7MUDI-425 | TTCCGGCATT | PlateE_A06 | i5MUDI-425 | AGGATTGTTC | GAACAATCCT
PlateE_B06 | i7MUDI-426 | CCACTTGTGG | PlateE_B06 | i5MUDI-426 | CTTGTGTCGG | CCGACACAAG
PlateE_C06 | i7MUDI-427 | TGGCTATTAC | PlateE_C06 | i5MUDI-427 | GCTTCTCTAA | TTAGAGAAGC
PlateE_D06 | i7MUDI-428 | TCCGCACACA | PlateE_D06 | i5MUDI-428 | TAGAACCTCT | AGAGGTTCTA
PlateE_E06 | i7MUDI-429 | GTAACACCGG | PlateE_E06 | i5MUDI-429 | AATAGCAGTG | CACTGCTATT
PlateE_F06 | i7MUDI-430 | TGTTATGGCT | PlateE_F06 | i5MUDI-430 | ACCAAGTAAG | CTTACTTGGT
PlateE_G06 | i7MUDI-431 | TTACAGGTTG | PlateE_G06 | i5MUDI-431 | CGAGAAGGTT | AACCTTCTCG
PlateE_H06 | i7MUDI-432 | GAATCTCAAC | PlateE_H06 | i5MUDI-432 | GAAGAGACGC | GCGTCTCTTC
PlateE_A07 | i7MUDI-433 | CGTGATGCGA | PlateE_A07 | i5MUDI-433 | TTCAACCACA | TGTGGTTGAA
PlateE_B07 | i7MUDI-434 | CTAACGAGCG | PlateE_B07 | i5MUDI-434 | AAGACTGACA | TGTCAGTCTT
PlateE_C07 | i7MUDI-435 | TTGTGTTCCT | PlateE_C07 | i5MUDI-435 | AGGCCGATAA | TTATCGGCCT
PlateE_D07 | i7MUDI-436 | GACGTTACGG | PlateE_D07 | i5MUDI-436 | CTTATACCTC | GAGGTATAAG
PlateE_E07 | i7MUDI-437 | ATCCAAGGAC | PlateE_E07 | i5MUDI-437 | GAGTCAACGT | ACGTTGACTC
PlateE_F07 | i7MUDI-438 | ATCGGCGCTA | PlateE_F07 | i5MUDI-438 | TCTCGGTTCG | CGAACCGAGA
PlateE_G07 | i7MUDI-439 | AGGATTGAAG | PlateE_G07 | i5MUDI-439 | AGCGCAACGA | TCGTTGCGCT
PlateE_H07 | i7MUDI-440 | ATCTACTCCT | PlateE_H07 | i5MUDI-440 | ATACTCACCT | AGGTGAGTAT
PlateE_A08 | i7MUDI-441 | CGTCACGTAG | PlateE_A08 | i5MUDI-441 | CGAGCTTGAC | GTCAAGCTCG
PlateE_B08 | i7MUDI-442 | TGTGTCGAGC | PlateE_B08 | i5MUDI-442 | GAATCCTTGA | TCAAGGATTC
PlateE_C08 | i7MUDI-443 | CCGCTCCTTA | PlateE_C08 | i5MUDI-443 | TCTGCCAATG | CATTGGCAGA
PlateE_D08 | i7MUDI-444 | TCACACTGAA | PlateE_D08 | i5MUDI-444 | ATCTGACCGA | TCGGTCAGAT
PlateE_E08 | i7MUDI-445 | AGTCGGCCAT | PlateE_E08 | i5MUDI-445 | AAGCTATGCA | TGCATAGCTT
PlateE_F08 | i7MUDI-446 | CGCCATTCTG | PlateE_F08 | i5MUDI-446 | CCTGGATGAG | CTCATCCAGG
PlateE_G08 | i7MUDI-447 | AACGACTGGC | PlateE_G08 | i5MUDI-447 | GTGATCTCTT | AAGAGATCAC
PlateE_H08 | i7MUDI-448 | GCTCAATACA | PlateE_H08 | i5MUDI-448 | TGCCTAACGC | GCGTTAGGCA
PlateE_A09 | i7MUDI-774 | TCGCCAACCA | PlateE_A09 | i5MUDI-774 | TGCCAAGGAG | CTCCTTGGCA
PlateE_B09 | i7MUDI-450 | AGAACACGAT | PlateE_B09 | i5MUDI-450 | ACTGTTCATC | GATGAACAGT
PlateE_C09 | i7MUDI-451 | CGGATAACCG | PlateE_C09 | i5MUDI-451 | CTCACAGGAA | TTCCTGTGAG
PlateE_D09 | i7MUDI-452 | TGTTCACCTC | PlateE_D09 | i5MUDI-452 | GTTCGCCAGT | ACTGGCGAAC
PlateE_E09 | i7MUDI-453 | AGGCGAGAGA | PlateE_E09 | i5MUDI-453 | TAACAGGCCG | CGGCCTGTTA
PlateE_F09 | i7MUDI-454 | CCGGTAATAA | PlateE_F09 | i5MUDI-454 | ACTCGAACGC | GCGTTCGAGT
PlateE_G09 | i7MUDI-455 | TACTTGTCCT | PlateE_G09 | i5MUDI-455 | ATACATTCGC | GCGAATGTAT
PlateE_H09 | i7MUDI-456 | TATTCGGCGG | PlateE_H09 | i5MUDI-456 | CCACGCATAA | TTATGCGTGG
PlateE_A10 | i7MUDI-457 | TTATGTGCAC | PlateE_A10 | i5MUDI-457 | GTACCTAGTG | CACTAGGTAC
PlateE_B10 | i7MUDI-458 | TCGTAGTGTA | PlateE_B10 | i5MUDI-458 | TGGACATCCT | AGGATGTCCA
PlateE_C10 | i7MUDI-459 | TTCGCCGAAC | PlateE_C10 | i5MUDI-459 | TCCAACTCTA | TAGAGTTGGA
PlateE_D10 | i7MUDI-460 | GAAGGTGCCT | PlateE_D10 | i5MUDI-460 | AGCGAGAGTG | CACTCTCGCT
PlateE_E10 | i7MUDI-461 | CAATTAGGAG | PlateE_E10 | i5MUDI-461 | CCGAACAAGT | ACTTGTTCGG
PlateE_F10 | i7MUDI-462 | AGGTTAGGTC | PlateE_F10 | i5MUDI-462 | TTGGCATTAC | GTAATGCCAA
PlateE_G10 | i7MUDI-463 | CAAGGCGTGA | PlateE_G10 | i5MUDI-463 | ACAACGTTCA | TGAACGTTGT
PlateE_H10 | i7MUDI-464 | TCCAAGAATG | PlateE_H10 | i5MUDI-464 | CGCATTAAGG | CCTTAATGCG
PlateE_A11 | i7MUDI-465 | ATTGCCTGTT | PlateE_A11 | i5MUDI-465 | TCCTTCAGGC | GCCTGAAGGA
PlateE_B11 | i7MUDI-775 | GAGGCTTGGT | PlateE_B11 | i5MUDI-775 | GTGAGGATGG | CCATCCTCAC
PlateE_C11 | i7MUDI-467 | ATTATGAGGC | PlateE_C11 | i5MUDI-467 | CCGCTAACAG | CTGTTAGCGG
PlateE_D11 | i7MUDI-468 | GAGACCATAA | PlateE_D11 | i5MUDI-468 | TTCGCTCGGA | TCCGAGCGAA
PlateE_E11 | i7MUDI-469 | TGGAGTACAA | PlateE_E11 | i5MUDI-469 | CTACGGCCTT | AAGGCCGTAG
PlateE_F11 | i7MUDI-470 | TCAAGGTCAT | PlateE_F11 | i5MUDI-470 | ACATAAGGTG | CACCTTATGT
PlateE_G11 | i7MUDI-471 | CTCGAATCGG | PlateE_G11 | i5MUDI-471 | CTAGGCGAGA | TCTCGCCTAG
PlateE_H11 | i7MUDI-472 | TCATATCGTC | PlateE_H11 | i5MUDI-472 | TGAACGAGCT | AGCTCGTTCA
PlateE_A12 | i7MUDI-473 | CCTTCTACAA | PlateE_A12 | i5MUDI-473 | AACACCTAGT | ACTAGGTGTT
PlateE_B12 | i7MUDI-474 | CAGGCACTTC | PlateE_B12 | i5MUDI-474 | AGATCATCAG | CTGATGATCT
PlateE_C12 | i7MUDI-475 | TCCATTCCAT | PlateE_C12 | i5MUDI-475 | TTCCTCCTGT | ACAGGAGGAA
PlateE_D12 | i7MUDI-476 | ATGTGGCTCG | PlateE_D12 | i5MUDI-476 | AATGAAGCTC | GAGCTTCATT
PlateE_E12 | i7MUDI-477 | GTATAGACTC | PlateE_E12 | i5MUDI-477 | CGCCGATTGT | ACAATCGGCG
PlateE_F12 | i7MUDI-478 | GACTAACCAA | PlateE_F12 | i5MUDI-478 | TCTCTGAAGA | TCTTCAGAGA
PlateE_G12 | i7MUDI-479 | CTGGTACCAT | PlateE_G12 | i5MUDI-479 | CATTCCGAGG | CCTCGGAATG
PlateE_H12 | i7MUDI-480 | TTGTCTGGCG | PlateE_H12 | i5MUDI-480 | TTCGAGTCAT | ATGACTCGAA
PlateF_A01 | i7MUDI-481 | GTTCGACAAT | PlateF_A01 | i5MUDI-481 | ATCTCCTCTT | AAGAGGAGAT
PlateF_B01 | i7MUDI-482 | TGGTAGGTGG | PlateF_B01 | i5MUDI-482 | CGACTCATGG | CCATGAGTCG
PlateF_C01 | i7MUDI-483 | GTAACCGATC | PlateF_C01 | i5MUDI-483 | GCGTATCGCA | TGCGATACGC
PlateF_D01 | i7MUDI-484 | CACCTCACCA | PlateF_D01 | i5MUDI-484 | TATGGAGAAC | GTTCTCCATA
PlateF_E01 | i7MUDI-485 | CCTGATTGTT | PlateF_E01 | i5MUDI-485 | AGCTCCAAGG | CCTTGGAGCT
PlateF_F01 | i7MUDI-486 | TGCACACCAG | PlateF_F01 | i5MUDI-486 | CCAACCGTCT | AGACGGTTGG
PlateF_G01 | i7MUDI-487 | TTCTCCACGC | PlateF_G01 | i5MUDI-487 | GACCTTCGTC | GACGAAGGTC
PlateF_H01 | i7MUDI-488 | ATCTTGCACA | PlateF_H01 | i5MUDI-488 | TTGGTCACAA | TTGTGACCAA
PlateF_A02 | i7MUDI-489 | AGCCTGTATT | PlateF_A02 | i5MUDI-489 | ATCGATGCTA | TAGCATCGAT
PlateF_B02 | i7MUDI-490 | GAAGGCAACG | PlateF_B02 | i5MUDI-490 | CGCAGATACG | CGTATCTGCG
PlateF_C02 | i7MUDI-491 | TGAATAGGAC | PlateF_C02 | i5MUDI-491 | GCATGTTAGT | ACTAACATGC
PlateF_D02 | i7MUDI-492 | CCGCAATGGA | PlateF_D02 | i5MUDI-492 | TAGCCTCAAC | GTTGAGGCTA
PlateF_E02 | i7MUDI-493 | CTATGAGGTT | PlateF_E02 | i5MUDI-493 | ATTGGTAAGG | CCTTACCAAT
PlateF_F02 | i7MUDI-494 | GCCGATATAG | PlateF_F02 | i5MUDI-494 | CCGAGAAGAC | GTCTTCTCGG
PlateF_G02 | i7MUDI-495 | GCTAGTCAGC | PlateF_G02 | i5MUDI-495 | GAGAGCTACT | AGTAGCTCTC
PlateF_H02 | i7MUDI-496 | CGGTCTAACA | PlateF_H02 | i5MUDI-496 | TGCCTGATTA | TAATCAGGCA
PlateF_A03 | i7MUDI-497 | TTCAATCGCT | PlateF_A03 | i5MUDI-497 | ATGCGCAATT | AATTGCGCAT
PlateF_B03 | i7MUDI-498 | TTGGCCAATG | PlateF_B03 | i5MUDI-498 | CGAGGATTAA | TTAATCCTCG
PlateF_C03 | i7MUDI-499 | AAGGATATGC | PlateF_C03 | i5MUDI-499 | GAATGGACCG | CGGTCCATTC
PlateF_D03 | i7MUDI-500 | CGAGGACCAA | PlateF_D03 | i5MUDI-500 | TCTGCGTTGC | GCAACGCAGA
PlateF_E03 | i7MUDI-501 | ACCATGGCCT | PlateF_E03 | i5MUDI-501 | ACAAGCTCCT | AGGAGCTTGT
PlateF_F03 | i7MUDI-502 | AATCACCAAG | PlateF_F03 | i5MUDI-502 | CAACATACAC | GTGTATGTTG
PlateF_G03 | i7MUDI-503 | GCATAAGCGC | PlateF_G03 | i5MUDI-503 | GTTGACAGGA | TCCTGTCAAC
PlateF_H03 | i7MUDI-504 | AGGAGGCGTA | PlateF_H03 | i5MUDI-504 | TGACAGCGTG | CACGCTGTCA
PlateF_A04 | i7MUDI-505 | CGTACTCATT | PlateF_A04 | i5MUDI-505 | AACAAGGTAC | GTACCTTGTT
PlateF_B04 | i7MUDI-506 | TAAGCGCGCT | PlateF_B04 | i5MUDI-506 | ACCACCGGTA | TACCGGTGGT
PlateF_C04 | i7MUDI-507 | AGACTACTTG | PlateF_C04 | i5MUDI-507 | CGTAACACCT | AGGTGTTACG
PlateF_D04 | i7MUDI-508 | TACGCAGTGC | PlateF_D04 | i5MUDI-508 | GTCAACCTAG | CTAGGTTGAC
PlateF_E04 | i7MUDI-509 | GCGCTTACAA | PlateF_E04 | i5MUDI-509 | TAATCTAGGC | GCCTAGATTA
PlateF_F04 | i7MUDI-776 | GAATTCGCGA | PlateF_F04 | i5MUDI-776 | TCCACATTGT | ACAATGTGGA
PlateF_G04 | i7MUDI-511 | CGTGTAGCCT | PlateF_G04 | i5MUDI-511 | ACCATTCCAA | TTGGAATGGT
PlateF_H04 | i7MUDI-512 | CCTCCTCTTG | PlateF_H04 | i5MUDI-512 | CGTTATTGTC | GACAATAACG
PlateF_A05 | i7MUDI-513 | AGTGTTCAAC | PlateF_A05 | i5MUDI-513 | GTCGGTCTGT | ACAGACCGAC
PlateF_B05 | i7MUDI-514 | CATAACAGGA | PlateF_B05 | i5MUDI-514 | TACGTATTCG | CGAATACGTA
PlateF_C05 | i7MUDI-515 | TATCGCACTT | PlateF_C05 | i5MUDI-515 | AACCAGCCAG | CTGGCTGGTT
PlateF_D05 | i7MUDI-516 | TCCTAAGGTT | PlateF_D05 | i5MUDI-516 | AGCGGACCAT | ATGGTCCGCT
PlateF_E05 | i7MUDI-517 | GTTATTGCGG | PlateF_E05 | i5MUDI-517 | CCGCTACTTA | TAAGTAGCGG
PlateF_F05 | i7MUDI-518 | TCGTTGGAAC | PlateF_F05 | i5MUDI-518 | GAACACCACG | CGTGGTGTTC
PlateF_G05 | i7MUDI-519 | TCCTCCGTCA | PlateF_G05 | i5MUDI-519 | TTAAGCCTGC | GCAGGCTTAA
PlateF_H05 | i7MUDI-520 | GTATGTCGCT | PlateF_H05 | i5MUDI-520 | CCTGTACTCG | CGAGTACAGG
PlateF_A06 | i7MUDI-521 | AAGGTCGATT | PlateF_A06 | i5MUDI-521 | AGGCAAGAGC | GCTCTTGCCT
PlateF_B06 | i7MUDI-522 | TCGGAAGACG | PlateF_B06 | i5MUDI-522 | CCTCAAGATT | AATCTTGAGG
PlateF_C06 | i7MUDI-523 | ACCGCCTTAC | PlateF_C06 | i5MUDI-523 | GACCGTTAAG | CTTAACGGTC
PlateF_D06 | i7MUDI-524 | TGTTCGTGGA | PlateF_D06 | i5MUDI-524 | TTACTGTCCA | TGGACAGTAA
PlateF_E06 | i7MUDI-525 | AGAGAGGTAG | PlateF_E06 | i5MUDI-525 | ACAGTGCCGA | TCGGCACTGT
PlateF_F06 | i7MUDI-526 | GCGACGTTCT | PlateF_F06 | i5MUDI-526 | AGGTGTGTGA | TCACACACCT
PlateF_G06 | i7MUDI-527 | GAGTTGCCAG | PlateF_G06 | i5MUDI-527 | CTAATTGCAG | CTGCAATTAG
PlateF_H06 | i7MUDI-528 | TTACCTTAGC | PlateF_H06 | i5MUDI-528 | GCTTCCATTC | GAATGGAAGC
PlateF_A07 | i7MUDI-529 | TACTCGGATA | PlateF_A07 | i5MUDI-529 | TAGTGTGGCT | AGCCACACTA
PlateF_B07 | i7MUDI-530 | CTACCGAATT | PlateF_B07 | i5MUDI-530 | CACAAGCTTA | TAAGCTTGTG
PlateF_C07 | i7MUDI-531 | TTGCGAGGCT | PlateF_C07 | i5MUDI-531 | AGTTCTACCA | TGGTAGAACT
PlateF_D07 | i7MUDI-532 | ATGTCTCCGG | PlateF_D07 | i5MUDI-532 | CTGTTCCGGT | ACCGGAACAG
PlateF_E07 | i7MUDI-533 | CAACAAGTTC | PlateF_E07 | i5MUDI-533 | GAGTGGCTTC | GAAGCCACTC
PlateF_F07 | i7MUDI-534 | CGACCTTGAA | PlateF_F07 | i5MUDI-534 | TCCTCTCGAG | CTCGAGAGGA
PlateF_G07 | i7MUDI-535 | GTCAGACCTA | PlateF_G07 | i5MUDI-535 | ATACGAGGAG | CTCCTCGTAT
PlateF_H07 | i7MUDI-536 | CCAGCTGCAT | PlateF_H07 | i5MUDI-536 | ATATTCGACG | CGTCGAATAT
PlateF_A08 | i7MUDI-777 | TGAGAATTGC | PlateF_A08 | i5MUDI-777 | TGGCCATATC | GATATGGCCA
PlateF_B08 | i7MUDI-538 | CAGTTATCTC | PlateF_B08 | i5MUDI-538 | GCTTAACTGT | ACAGTTAAGC
PlateF_C08 | i7MUDI-539 | TACCACGCGA | PlateF_C08 | i5MUDI-539 | TGTGTCCTAC | GTAGGACACA
PlateF_D08 | i7MUDI-540 | CTTCGAGGAG | PlateF_D08 | i5MUDI-540 | CCGATGGTTG | CAACCATCGG
PlateF_E08 | i7MUDI-541 | TACGCCGGTT | PlateF_E08 | i5MUDI-541 | AACAGAATGC | GCATTCTGTT
PlateF_F08 | i7MUDI-542 | ATCAGATAGG | PlateF_F08 | i5MUDI-542 | CCTGAGCGAA | TTCGCTCAGG
PlateF_G08 | i7MUDI-543 | TCTTCTTGAC | PlateF_G08 | i5MUDI-543 | GTAACGGATT | AATCCGTTAC
PlateF_H08 | i7MUDI-544 | ACTACTAGCA | PlateF_H08 | i5MUDI-544 | TGTTGTAGAG | CTCTACAACA
PlateF_A09 | i7MUDI-545 | TAGACCAACT | PlateF_A09 | i5MUDI-545 | ACCAACATGA | TCATGTTGGT
PlateF_B09 | i7MUDI-546 | AATGCTTCCT | PlateF_B09 | i5MUDI-546 | ACACTGTAGC | GCTACAGTGT
PlateF_C09 | i7MUDI-547 | CCAACTAAGG | PlateF_C09 | i5MUDI-547 | CTCCGCTCAA | TTGAGCGGAG
PlateF_D09 | i7MUDI-548 | ACACCGGAAC | PlateF_D09 | i5MUDI-548 | GATACTATGG | CCATAGTATC
PlateF_E09 | i7MUDI-549 | AACCAGGTTA | PlateF_E09 | i5MUDI-549 | TGGCTCTTCT | AGAAGAGCCA
PlateF_F09 | i7MUDI-550 | GCACGAACTT | PlateF_F09 | i5MUDI-550 | AATCCACTAC | GTAGTGGATT
PlateF_G09 | i7MUDI-551 | GTGGATTCAT | PlateF_G09 | i5MUDI-551 | AGAAGAGGTC | GACCTCTTCT
PlateF_H09 | i7MUDI-552 | AATTAAGCCG | PlateF_H09 | i5MUDI-552 | CCATAGGCGT | ACGCCTATGG
PlateF_A10 | i7MUDI-553 | GATTCATAGC | PlateF_A10 | i5MUDI-553 | GAACTCGGAA | TTCCGAGTTC
PlateF_B10 | i7MUDI-554 | GTCCATCTTA | PlateF_B10 | i5MUDI-554 | TTCAGGAATG | CATTCCTGAA
PlateF_C10 | i7MUDI-555 | CGCAAGGTTG | PlateF_C10 | i5MUDI-555 | AAGACGATTG | CAATCGTCTT
PlateF_D10 | i7MUDI-556 | TCCGTGACTT | PlateF_D10 | i5MUDI-556 | AGAGGCACTC | GAGTGCCTCT
PlateF_E10 | i7MUDI-557 | CTTATGGACG | PlateF_E10 | i5MUDI-557 | CTTCTGACAT | ATGTCAGAAG
PlateF_F10 | i7MUDI-558 | GACAGTTGTC | PlateF_F10 | i5MUDI-558 | GAGTAAGGAG | CTCCTTACTC
PlateF_G10 | i7MUDI-559 | CTCGTGTTAA | PlateF_G10 | i5MUDI-559 | TCTCAAGCGA | TCGCTTGAGA
PlateF_H10 | i7MUDI-560 | TCTGCGGTAT | PlateF_H10 | i5MUDI-560 | AAGAGGTCTT | AAGACCTCTT
PlateF_A11 | i7MUDI-561 | TCTCTCGATT | PlateF_A11 | i5MUDI-561 | AATCATGTGG | CCACATGATT
PlateF_B11 | i7MUDI-562 | TTGAGTATCG | PlateF_B11 | i5MUDI-562 | CGACTAACTA | TAGTTAGTCG
PlateF_C11 | i7MUDI-563 | CCTGCAACTC | PlateF_C11 | i5MUDI-563 | TCACGGTGAT | ATCACCGTGA
PlateF_D11 | i7MUDI-564 | TATGCCACAA | PlateF_D11 | i5MUDI-564 | ACAGCCAACA | TGTTGGCTGT
PlateF_E11 | i7MUDI-565 | CTGACACGCT | PlateF_E11 | i5MUDI-565 | AAGGCTAGTT | AACTAGCCTT
PlateF_F11 | i7MUDI-566 | GATCGGTATT | PlateF_F11 | i5MUDI-566 | TCACACAGTA | TACTGTGTGA
PlateF_G11 | i7MUDI-567 | TCTAGAGGCG | PlateF_G11 | i5MUDI-567 | ACATAGTGAC | GTCACTATGT
PlateF_H11 | i7MUDI-568 | AATCTCTGTC | PlateF_H11 | i5MUDI-568 | CGGTTCAGTG | CACTGAACCG
PlateF_A12 | i7MUDI-569 | GTTAGCGTAA | PlateF_A12 | i5MUDI-569 | TAGAGAGTGT | ACACTCTCTA
PlateF_B12 | i7MUDI-570 | GTCGACGGAA | PlateF_B12 | i5MUDI-570 | TCTTGCACGG | CCGTGCAAGA
PlateF_C12 | i7MUDI-571 | CTACGGCGAT | PlateF_C12 | i5MUDI-571 | CATTGCATAG | CTATGCAATG
PlateF_D12 | i7MUDI-572 | CAGTAGAATG | PlateF_D12 | i5MUDI-572 | TGGTACGCTT | AAGCGTACCA
PlateF_E12 | i7MUDI-573 | TTGCTGCTTC | PlateF_E12 | i5MUDI-573 | AGGATCAGGC | GCCTGATCCT
PlateF_F12 | i7MUDI-574 | GTACAGGACA | PlateF_F12 | i5MUDI-574 | CTTAGCGGAG | CTCCGCTAAG
PlateF_G12 | i7MUDI-575 | CATAGTGGCT | PlateF_G12 | i5MUDI-575 | AACACACTCA | TGAGTGTGTT
PlateF_H12 | i7MUDI-576 | CTAGTCCGTG | PlateF_H12 | i5MUDI-576 | CATGCATCTT | AAGATGCATG
PlateG_A01 | i7MUDI-577 | GTCCGGTTAT | PlateG_A01 | i5MUDI-577 | AAGCTCTCTG | CAGAGAGCTT
PlateG_B01 | i7MUDI-778 | CGGCTTAGGA | PlateG_B01 | i5MUDI-778 | CGTGAGAATT | AATTCTCACG
PlateG_C01 | i7MUDI-579 | CACTTCAGGC | PlateG_C01 | i5MUDI-579 | GTTGTCGCGT | ACGCGACAAC
PlateG_D01 | i7MUDI-580 | ACGGCATATA | PlateG_D01 | i5MUDI-580 | TCTGATCTCA | TGAGATCAGA
PlateG_E01 | i7MUDI-581 | TCAACTGGTT | PlateG_E01 | i5MUDI-581 | ACCACCACAC | GTGTGGTGGT
PlateG_F01 | i7MUDI-582 | CTCAGTTGCG | PlateG_F01 | i5MUDI-582 | CAAGGTTACT | AGTAACCTTG
PlateG_G01 | i7MUDI-583 | TTGTTACAGC | PlateG_G01 | i5MUDI-583 | GTGTTACTGA | TCAGTAACAC
PlateG_H01 | i7MUDI-584 | ACCAACAGAA | PlateG_H01 | i5MUDI-584 | TGGAGTCTTG | CAAGACTCCA
PlateG_A02 | i7MUDI-585 | AGCGGACGTT | PlateG_A02 | i5MUDI-585 | AATCCTAACG | CGTTAGGATT
PlateG_B02 | i7MUDI-586 | TAACTACACG | PlateG_B02 | i5MUDI-586 | CCTTGGAAGT | ACTTCCAAGG
PlateG_C02 | i7MUDI-587 | CTATTCGAAC | PlateG_C02 | i5MUDI-587 | GTAAGACAAC | GTTGTCTTAC
PlateG_D02 | i7MUDI-588 | ATTAGTCCGA | PlateG_D02 | i5MUDI-588 | TGCATAGCTA | TAGCTATGCA
PlateG_E02 | i7MUDI-589 | TCGAGAAGTT | PlateG_E02 | i5MUDI-589 | ACCTGTGCCT | AGGCACAGGT
PlateG_F02 | i7MUDI-590 | GACATGTGCG | PlateG_F02 | i5MUDI-590 | CTAACAATGC | GCATTGTTAG
PlateG_G02 | i7MUDI-591 | GAGAGCGAAC | PlateG_G02 | i5MUDI-591 | GAACGACTTA | TAAGTCGTTC
PlateG_H02 | i7MUDI-592 | CGATGTTAGA | PlateG_H02 | i5MUDI-592 | TGGATAAGAG | CTCTTATCCA
PlateG_A03 | i7MUDI-593 | TTCCTCTCTT | PlateG_A03 | i5MUDI-593 | AGAGTATCCT | AGGATACTCT
PlateG_B03 | i7MUDI-594 | AAGGAACCGG | PlateG_B03 | i5MUDI-594 | CTGGAGGATA | TATCCTCCAG
PlateG_C03 | i7MUDI-595 | GCAATACTAC | PlateG_C03 | i5MUDI-595 | GCTGGTCAAG | CTTGACCAGC
PlateG_D03 | i7MUDI-596 | ATAGCAGTCA | PlateG_D03 | i5MUDI-596 | TACGAGCTGC | GCAGCTCGTA
PlateG_E03 | i7MUDI-597 | CCAAGCCTCT | PlateG_E03 | i5MUDI-597 | ATGCCTGCTC | GAGCAGGCAT
PlateG_F03 | i7MUDI-598 | AATTGGACTG | PlateG_F03 | i5MUDI-598 | CGATCTGTCG | CGACAGATCG
PlateG_G03 | i7MUDI-599 | TCAGGTGTGC | PlateG_G03 | i5MUDI-599 | GAAGCGTGAT | ATCACGCTTC
PlateG_H03 | i7MUDI-600 | TCGCCTCTAA | PlateG_H03 | i5MUDI-600 | TCAATCCGGA | TCCGGATTGA
PlateG_A04 | i7MUDI-601 | AGTCCTTGCG | PlateG_A04 | i5MUDI-601 | ACCGCTGTGA | TCACAGCGGT
PlateG_B04 | i7MUDI-602 | CCGTTCTCCT | PlateG_B04 | i5MUDI-602 | ACATAACAGC | GCTGTTATGT
PlateG_C04 | i7MUDI-603 | CAAGAGTAGG | PlateG_C04 | i5MUDI-603 | CGCACTTCAT | ATGAAGTGCG
PlateG_D04 | i7MUDI-604 | GTTCAATGTC | PlateG_D04 | i5MUDI-604 | GAAGAATGTG | CACATTCTTC
PlateG_E04 | i7MUDI-605 | ACGGTGCCAA | PlateG_E04 | i5MUDI-605 | TTGGCAAGCA | TGCTTGCCAA
PlateG_F04 | i7MUDI-606 | AACACTGTTG | PlateG_F04 | i5MUDI-606 | ACCGCCTATC | GATAGGCGGT
PlateG_G04 | i7MUDI-607 | GTCATCAACT | PlateG_G04 | i5MUDI-607 | AGGTGCCGTT | AACGGCACCT
PlateG_H04 | i7MUDI-608 | TGGCGTAAGG | PlateG_H04 | i5MUDI-608 | CTAAGCTCGG | CCGAGCTTAG
PlateG_A05 | i7MUDI-609 | TCTACGCAAC | PlateG_A05 | i5MUDI-609 | GACCAGAGCA | TGCTCTGGTC
PlateG_B05 | i7MUDI-610 | CTGCAGCGTA | PlateG_B05 | i5MUDI-610 | TCCTGGCTAC | GTAGCCAGGA
PlateG_C05 | i7MUDI-611 | GAGGTGGTTG | PlateG_C05 | i5MUDI-611 | ACTTCCAGGA | TCCTGGAAGT
PlateG_D05 | i7MUDI-612 | GTGTAAGGAT | PlateG_D05 | i5MUDI-612 | AGTTACGAAG | CTTCGTAACT
PlateG_E05 | i7MUDI-613 | CAGAACGTCG | PlateG_E05 | i5MUDI-613 | CCGTGAACCA | TGGTTCACGG
PlateG_F05 | i7MUDI-614 | TAACGCCAGC | PlateG_F05 | i5MUDI-614 | GTGAAGTATC | GATACTTCAC
PlateG_G05 | i7MUDI-615 | GACGTCCGTA | PlateG_G05 | i5MUDI-615 | TAACGCAAGT | ACTTGCGTTA
PlateG_H05 | i7MUDI-616 | GTTCGGCTTG | PlateG_H05 | i5MUDI-616 | CCTACACGGT | ACCGTGTAGG
PlateG_A06 | i7MUDI-617 | GTGGCGAGAT | PlateG_A06 | i5MUDI-617 | ATCTGGTGAA | TTCACCAGAT
PlateG_B06 | i7MUDI-779 | AGCCATCCAC | PlateG_B06 | i5MUDI-779 | GTATTAGGAC | GTCCTAATAC
PlateG_C06 | i7MUDI-780 | GTCCTAATAC | PlateG_C06 | i5MUDI-780 | GTGGATGGCT | AGCCATCCAC
PlateG_D06 | i7MUDI-620 | GAATGGTGTA | PlateG_D06 | i5MUDI-620 | TACGTCGGTG | CACCGACGTA
PlateG_E06 | i7MUDI-621 | GATCTAGGCG | PlateG_E06 | i5MUDI-621 | ATGAGGAACA | TGTTCCTCAT
PlateG_F06 | i7MUDI-622 | CAGGAAGGCT | PlateG_F06 | i5MUDI-622 | ATGCGGTCGA | TCGACCGCAT
PlateG_G06 | i7MUDI-623 | CAGGTCACAG | PlateG_G06 | i5MUDI-623 | CGCCAACCTT | AAGGTTGGCG
PlateG_H06 | i7MUDI-624 | AGATGCCTTC | PlateG_H06 | i5MUDI-624 | GCAGTGATCG | CGATCACTGC
PlateG_A07 | i7MUDI-625 | TGGTAATCGA | PlateG_A07 | i5MUDI-625 | TAGTTGCCAC | GTGGCAACTA
PlateG_B07 | i7MUDI-626 | CTCTTGTGTT | PlateG_B07 | i5MUDI-626 | ATTCAGCCTA | TAGGCTGAAT
PlateG_C07 | i7MUDI-627 | CTTCATGCAT | PlateG_C07 | i5MUDI-627 | AATCGTTGTC | GACAACGATT
PlateG_D07 | i7MUDI-628 | ACCGATTGCG | PlateG_D07 | i5MUDI-628 | CGATCATACT | AGTATGATCG
PlateG_E07 | i7MUDI-629 | ACCTATGTTC | PlateG_E07 | i5MUDI-629 | GCTTATGAGA | TCTCATAAGC
PlateG_F07 | i7MUDI-630 | GCAGTAGAGA | PlateG_F07 | i5MUDI-630 | TTGTGTGTAG | CTACACACAA
PlateG_G07 | i7MUDI-631 | TGTACCGTCG | PlateG_G07 | i5MUDI-631 | ACAGGAACCG | CGGTTCCTGT
PlateG_H07 | i7MUDI-632 | CATCCAGATT | PlateG_H07 | i5MUDI-632 | AGGACTCCGA | TCGGAGTCCT
PlateG_A08 | i7MUDI-633 | TTCTTGATCG | PlateG_A08 | i5MUDI-633 | CCGTCTGAAC | GTTCAGACGG
PlateG_B08 | i7MUDI-634 | ATTGCGGAGC | PlateG_B08 | i5MUDI-634 | GTCCAGACTG | CAGTCTGGAC
PlateG_C08 | i7MUDI-635 | CTCTGGACAA | PlateG_C08 | i5MUDI-635 | TATCAGCACT | AGTGCTGATA
PlateG_D08 | i7MUDI-636 | GTGTCCTATT | PlateG_D08 | i5MUDI-636 | ACCTGGAATC | GATTCCAGGT
PlateG_E08 | i7MUDI-637 | ACACCGTCCT | PlateG_E08 | i5MUDI-637 | AATGAACGAG | CTCGTTCATT
PlateG_F08 | i7MUDI-638 | TGAGAGCCTG | PlateG_F08 | i5MUDI-638 | CGAGAACAGA | TCTGTTCTCG
PlateG_G08 | i7MUDI-639 | TATATCTCGC | PlateG_G08 | i5MUDI-639 | GCTGTAGGTC | GACCTACAGC
PlateG_H08 | i7MUDI-640 | AGTGTCGTAA | PlateG_H08 | i5MUDI-640 | TTCCTGAGCT | AGCTCAGGAA
PlateG_A09 | i7MUDI-641 | ATGGTTCTTG | PlateG_A09 | i5MUDI-641 | ATGAGTGGTG | CACCACTCAT
PlateG_B09 | i7MUDI-642 | CGCCTAAGCT | PlateG_B09 | i5MUDI-642 | ACTTCCTACT | AGTAGGAAGT
PlateG_C09 | i7MUDI-643 | CGCACCAATG | PlateG_C09 | i5MUDI-643 | CGTAATGGAA | TTCCATTACG
PlateG_D09 | i7MUDI-644 | AACGTAGAAC | PlateG_D09 | i5MUDI-644 | GAATATTCGG | CCGAATATTC
PlateG_E09 | i7MUDI-645 | ATCCGCTGGA | PlateG_E09 | i5MUDI-645 | TCAAGCAATC | GATTGCTTGA
PlateG_F09 | i7MUDI-646 | CCTGACGATG | PlateG_F09 | i5MUDI-646 | CCTCTCAGCA | TGCTGAGAGG
PlateG_G09 | i7MUDI-647 | GCATCGCCAT | PlateG_G09 | i5MUDI-647 | ACCGAAGGAT | ATCCTTCGGT
PlateG_H09 | i7MUDI-648 | GCTCCTTAGG | PlateG_H09 | i5MUDI-648 | CATTCGCGTC | GACGCGAATG
PlateG_A10 | i7MUDI-649 | AAGGCTGCTC | PlateG_A10 | i5MUDI-649 | GTTACCGCTA | TAGCGGTAAC
PlateG_B10 | i7MUDI-650 | GAGCCTGACA | PlateG_B10 | i5MUDI-650 | TTCCGATAGG | CCTATCGGAA
PlateG_C10 | i7MUDI-651 | TAGTCGTGAG | PlateG_C10 | i5MUDI-651 | AGAAGGTATG | CATACCTTCT
PlateG_D10 | i7MUDI-652 | CTGGCCTTAT | PlateG_D10 | i5MUDI-652 | AAGGTGAGAG | CTCTCACCTT
PlateG_E10 | i7MUDI-653 | GTCTAACAGG | PlateG_E10 | i5MUDI-653 | CAACTGCTGC | GCAGCAGTTG
PlateG_F10 | i7MUDI-654 | GTACGTATTC | PlateG_F10 | i5MUDI-654 | GTCCTCCATA | TATGGAGGAC
PlateG_G10 | i7MUDI-655 | CAGACATGTA | PlateG_G10 | i5MUDI-655 | TCTGAGCCTT | AAGGCTCAGA
PlateG_H10 | i7MUDI-656 | TGCCACCGTT | PlateG_H10 | i5MUDI-656 | CCGAATACTG | CAGTATTCGG
PlateG_A11 | i7MUDI-657 | GCCACCTAAT | PlateG_A11 | i5MUDI-657 | AATAGGCTGA | TCAGCCTATT
PlateG_B11 | i7MUDI-658 | CGTTGCATGG | PlateG_B11 | i5MUDI-658 | CTGTGGACAC | GTGTCCACAG
PlateG_C11 | i7MUDI-659 | TGGATTGTTC | PlateG_C11 | i5MUDI-659 | GCGTACGTTG | CAACGTACGC
PlateG_D11 | i7MUDI-660 | CCTCTAAGTA | PlateG_D11 | i5MUDI-660 | TCGTTACAGT | ACTGTAACGA
PlateG_E11 | i7MUDI-661 | TCGCAACATT | PlateG_E11 | i5MUDI-661 | ATTGATCTGC | GCAGATCAAT
PlateG_F11 | i7MUDI-662 | TACAGGCAGG | PlateG_F11 | i5MUDI-662 | ATCTACTTGG | CCAAGTAGAT
PlateG_G11 | i7MUDI-663 | TGTGGACTTA | PlateG_G11 | i5MUDI-663 | CCGCGATGTT | AACATCGCGG
PlateG_H11 | i7MUDI-664 | GATCATGGTT | PlateG_H11 | i5MUDI-664 | TAGAACGGTA | TACCGTTCTA
PlateG_A12 | i7MUDI-665 | GAAGATCTCG | PlateG_A12 | i5MUDI-665 | AACAAGACGT | ACGTCTTGTT
PlateG_B12 | i7MUDI-666 | GATCACTCTA | PlateG_B12 | i5MUDI-666 | CTTGCCACCA | TGGTGGCAAG
PlateG_C12 | i7MUDI-667 | GCTGCGTCTT | PlateG_C12 | i5MUDI-667 | CTACCGGTTC | GAACCGGTAG
PlateG_D12 | i7MUDI-668 | AGCTTACCGG | PlateG_D12 | i5MUDI-668 | ACGACAGAGG | CCTCTGTCGT
PlateG_E12 | i7MUDI-669 | TTGTTCGGTT | PlateG_E12 | i5MUDI-669 | CAAGTGACAA | TTGTCACTTG
PlateG_F12 | i7MUDI-670 | ATCTATCGAG | PlateG_F12 | i5MUDI-670 | ACAGTCGTAA | TTACGACTGT
PlateG_G12 | i7MUDI-671 | GCCGTTGGTT | PlateG_G12 | i5MUDI-671 | CTACAACATG | CATGTTGTAG
PlateG_H12 | i7MUDI-672 | TACAGATGAG | PlateG_H12 | i5MUDI-672 | ACCTTAACGG | CCGTTAAGGT
PlateH_A01 | i7MUDI-673 | CGCTCATAAT | PlateH_A01 | i5MUDI-673 | AATTAGCCGC | GCGGCTAATT
PlateH_B01 | i7MUDI-781 | AATTCTCACG | PlateH_B01 | i5MUDI-781 | TCCTAAGCCG | CGGCTTAGGA
PlateH_C01 | i7MUDI-675 | ACATGCACGC | PlateH_C01 | i5MUDI-675 | GCTTGCTGCA | TGCAGCAAGC
PlateH_D01 | i7MUDI-676 | GAGCCGCTTA | PlateH_D01 | i5MUDI-676 | TTGCCTACAT | ATGTAGGCAA
PlateH_E01 | i7MUDI-677 | ATCAGCGAAT | PlateH_E01 | i5MUDI-677 | ACATCTCTTC | GAAGAGATGT
PlateH_F01 | i7MUDI-678 | TATTGTTGCG | PlateH_F01 | i5MUDI-678 | CGCCAATGCA | TGCATTGGCG
PlateH_G01 | i7MUDI-679 | AGCGAGATTC | PlateH_G01 | i5MUDI-679 | GTACACTCAT | ATGAGTGTAC
PlateH_H01 | i7MUDI-680 | TGACGGACGA | PlateH_H01 | i5MUDI-680 | TAGCTGGTGG | CCACCAGCTA
PlateH_A02 | i7MUDI-681 | AAGGAGTGTT | PlateH_A02 | i5MUDI-681 | AATGCTTAGC | GCTAAGCATT
PlateH_B02 | i7MUDI-682 | ACATAGGAGG | PlateH_B02 | i5MUDI-682 | CGTCGGATCT | AGATCCGACG
PlateH_C02 | i7MUDI-683 | CTGCCTCAAC | PlateH_C02 | i5MUDI-683 | GCCGTGCTAA | TTAGCACGGC
PlateH_D02 | i7MUDI-684 | TGCCTTCACA | PlateH_D02 | i5MUDI-684 | TTCGCATGTG | CACATGCGAA
PlateH_E02 | i7MUDI-685 | CGAAGCTGTT | PlateH_E02 | i5MUDI-685 | AGAACCACCG | CGGTGGTTCT
PlateH_F02 | i7MUDI-686 | GTACTCGAGG | PlateH_F02 | i5MUDI-686 | CTGAGTTAGA | TCTAACTCAG
PlateH_G02 | i7MUDI-687 | GAGCGTCTAC | PlateH_G02 | i5MUDI-687 | GCTTAGTCAT | ATGACTAAGC
PlateH_H02 | i7MUDI-688 | CATGTGAACA | PlateH_H02 | i5MUDI-688 | TAGTGCAGTC | GACTGCACTA
PlateH_A03 | i7MUDI-689 | CTGAATCCTT | PlateH_A03 | i5MUDI-689 | AGTGTCTAGT | ACTAGACACT
PlateH_B03 | i7MUDI-690 | TAAGCACTAG | PlateH_B03 | i5MUDI-690 | CTACTCAGAC | GTCTGAGTAG
PlateH_C03 | i7MUDI-691 | GTCGGTAAGC | PlateH_C03 | i5MUDI-691 | GCTCAGCTTG | CAAGCTGAGC
PlateH_D03 | i7MUDI-692 | TAGGTCGCCA | PlateH_D03 | i5MUDI-692 | TAACGCGTCA | TGACGCGTTA
PlateH_E03 | i7MUDI-693 | GCCGAGAACT | PlateH_E03 | i5MUDI-693 | ACCTCTTCGC | GCGAAGAGGT
PlateH_F03 | i7MUDI-694 | TTAGGTTCGG | PlateH_F03 | i5MUDI-694 | CGTCCTGCTT | AAGCAGGACG
PlateH_G03 | i7MUDI-695 | GTAGAAGTAC | PlateH_G03 | i5MUDI-695 | GTCACAATCG | CGATTGTGAC
PlateH_H03 | i7MUDI-696 | TCTAACCGTA | PlateH_H03 | i5MUDI-696 | TAGACACCAA | TTGGTGTCTA
PlateH_A04 | i7MUDI-697 | CTAATGTCGG | PlateH_A04 | i5MUDI-697 | ACAGGACATA | TATGTCCTGT
PlateH_B04 | i7MUDI-698 | ACCTCCACAT | PlateH_B04 | i5MUDI-698 | ACGAATTGGC | GCCAATTCGT
PlateH_C04 | i7MUDI-699 | TGAGCCATGG | PlateH_C04 | i5MUDI-699 | CGCATGAGTT | AACTCATGCG
PlateH_D04 | i7MUDI-700 | GAACTGCGTC | PlateH_D04 | i5MUDI-700 | GTGTCGATCA | TGATCGACAC
PlateH_E04 | i7MUDI-701 | TTGCCGGTCA | PlateH_E04 | i5MUDI-701 | TACTACACAG | CTGTGTAGTA
PlateH_F04 | i7MUDI-702 | GATGGATCGG | PlateH_F04 | i5MUDI-702 | ACAGTGGACT | AGTCCACTGT
PlateH_G04 | i7MUDI-703 | AACCAACGCT | PlateH_G04 | i5MUDI-703 | ACTCATGGAC | GTCCATGAGT
PlateH_H04 | i7MUDI-704 | TGTTCGAATG | PlateH_H04 | i5MUDI-704 | CGATCGCCTA | TAGGCGATCG
PlateH_A05 | i7MUDI-705 | CCACAACAGC | PlateH_A05 | i5MUDI-705 | GTATACAGCG | CGCTGTATAC
PlateH_B05 | i7MUDI-706 | ACCTGTTGAA | PlateH_B05 | i5MUDI-706 | TACTCGCGGT | ACCGCGAGTA
PlateH_C05 | i7MUDI-707 | ACTGAGCTGG | PlateH_C05 | i5MUDI-707 | ACCTTGTACA | TGTACAAGGT
PlateH_D05 | i7MUDI-708 | TCTCCATGTT | PlateH_D05 | i5MUDI-708 | AGCTCTCACT | AGTGAGAGCT
PlateH_E05 | i7MUDI-709 | TTATAACGCG | PlateH_E05 | i5MUDI-709 | CCGCATGCAA | TTGCATGCGG
PlateH_F05 | i7MUDI-710 | CTGTAGGAGC | PlateH_F05 | i5MUDI-710 | GTTCATCCGG | CCGGATGAAC
PlateH_G05 | i7MUDI-711 | CTGGATCTAA | PlateH_G05 | i5MUDI-711 | TACTATCCTC | GAGGATAGTA
PlateH_H05 | i7MUDI-712 | ATGGACGAGG | PlateH_H05 | i5MUDI-712 | ATTAACGGCA | TGCCGTTAAT
PlateH_A06 | i7MUDI-713 | ATACGCTCAT | PlateH_A06 | i5MUDI-713 | AGCAGTTAAC | GTTAACTGCT
PlateH_B06 | i7MUDI-714 | TGCCTCTTGG | PlateH_B06 | i5MUDI-714 | CAAGGTCCGA | TCGGACCTTG
PlateH_C06 | i7MUDI-715 | TAGCGACCTC | PlateH_C06 | i5MUDI-715 | GTTGTCTATG | CATAGACAAC
PlateH_D06 | i7MUDI-716 | CCGAGTTCCA | PlateH_D06 | i5MUDI-716 | TACGCGATAT | ATATCGCGTA
PlateH_E06 | i7MUDI-717 | TGTGTGTCGG | PlateH_E06 | i5MUDI-717 | ACCACATCCG | CGGATGTGGT
PlateH_F06 | i7MUDI-718 | CGGTTGGCAT | PlateH_F06 | i5MUDI-718 | ACCTAGGTTA | TAACCTAGGT
PlateH_G06 | i7MUDI-719 | AGGAACAACG | PlateH_G06 | i5MUDI-719 | CGCGTACACT | AGTGTACGCG
PlateH_H06 | i7MUDI-720 | GTCTTCGTTC | PlateH_H06 | i5MUDI-720 | GATGGATAGG | CCTATCCATC
PlateH_A07 | i7MUDI-721 | ACGTCCTCGA | PlateH_A07 | i5MUDI-721 | TGTGAGCAAG | CTTGCTCACA
PlateH_B07 | i7MUDI-722 | ATTGGTGTGG | PlateH_B07 | i5MUDI-722 | AATGCCACGT | ACGTGGCATT
PlateH_C07 | i7MUDI-723 | GCGTTACCTT | PlateH_C07 | i5MUDI-723 | AGGAGGCTAC | GTAGCCTCCT
PlateH_D07 | i7MUDI-724 | GAATCGACCG | PlateH_D07 | i5MUDI-724 | CCGAGGTAAT | ATTACCTCGG
PlateH_E07 | i7MUDI-725 | GCGACTAGAC | PlateH_E07 | i5MUDI-725 | GACACGTATA | TATACGTGTC
PlateH_F07 | i7MUDI-726 | ATGCTACTGA | PlateH_F07 | i5MUDI-726 | TACGGCAACG | CGTTGCCGTA
PlateH_G07 | i7MUDI-727 | TCCACGTTAG | PlateH_G07 | i5MUDI-727 | ACGAACCGTG | CACGGTTCGT
PlateH_H07 | i7MUDI-728 | TGTCGCTTCT | PlateH_H07 | i5MUDI-728 | ACATCTGGCA | TGCCAGATGT
PlateH_A08 | i7MUDI-729 | AGTAAGGCGG | PlateH_A08 | i5MUDI-729 | CGTACCGAAT | ATTCGGTACG
PlateH_B08 | i7MUDI-730 | ACTGTACTTC | PlateH_B08 | i5MUDI-730 | TATCCTTCGG | CCGAAGGATA
PlateH_C08 | i7MUDI-731 | CCTAAGCTAA | PlateH_C08 | i5MUDI-731 | ACATCACGAT | ATCGTGATGT
PlateH_D08 | i7MUDI-732 | TCCTAGTCGG | PlateH_D08 | i5MUDI-732 | ACCAGAACTT | AAGTTCTGGT
PlateH_E08 | i7MUDI-733 | CCGGTGTATT | PlateH_E08 | i5MUDI-733 | ACCGACTCGT | ACGAGTCGGT
PlateH_F08 | i7MUDI-734 | CAGCATAGCG | PlateH_F08 | i5MUDI-734 | CAGTATCAAG | CTTGATACTG
PlateH_G08 | i7MUDI-735 | CCTCAGTGAC | PlateH_G08 | i5MUDI-735 | ACGTTCTGTA | TACAGAACGT
PlateH_H08 | i7MUDI-736 | TTACTGCAGA | PlateH_H08 | i5MUDI-736 | CGGATTGCGT | ACGCAATCCG
PlateH_A09 | i7MUDI-737 | AGAGACTCGG | PlateH_A09 | i5MUDI-737 | CGTGCTGAGA | TCTCAGCACG
PlateH_B09 | i7MUDI-738 | GTGATTGGTA | PlateH_B09 | i5MUDI-738 | ATTACGTCTG | CAGACGTAAT
PlateH_C09 | i7MUDI-739 | AGGTCATCCT | PlateH_C09 | i5MUDI-739 | CGGTCTCTAT | ATAGAGACCG
PlateH_D09 | i7MUDI-740 | ACGAATTAGG | PlateH_D09 | i5MUDI-740 | ACCTGACAAG | CTTGTCAGGT
PlateH_E09 | i7MUDI-741 | CATACGGTTA | PlateH_E09 | i5MUDI-741 | CGACTGGAGT | ACTCCAGTCG
PlateH_F09 | i7MUDI-742 | TGCTATCATG | PlateH_F09 | i5MUDI-742 | ACTTATCCAG | CTGGATAAGT
PlateH_G09 | i7MUDI-743 | CGATATTCCT | PlateH_G09 | i5MUDI-743 | ATGATCGCGG | CCGCGATCAT
PlateH_H09 | i7MUDI-744 | TGCGGTCTAG | PlateH_H09 | i5MUDI-744 | AATGGTCGGT | ACCGACCATT
PlateH_A10 | i7MUDI-745 | ACAAGCCATA | PlateH_A10 | i5MUDI-745 | AATCCGTCAT | ATGACGGATT
PlateH_B10 | i7MUDI-746 | TATGAGCCAT | PlateH_B10 | i5MUDI-746 | AGCCGATCTC | GAGATCGGCT
PlateH_C10 | i7MUDI-747 | CCATATCTAG | PlateH_C10 | i5MUDI-747 | CGATACCTGT | ACAGGTATCG
PlateH_D10 | i7MUDI-748 | GATTATCCTG | PlateH_D10 | i5MUDI-748 | CAGTAAGACA | TGTCTTACTG
PlateH_E10 | i7MUDI-749 | ACGTATCACA | PlateH_E10 | i5MUDI-749 | TGTGCTCTGG | CCAGAGCACA
PlateH_F10 | i7MUDI-750 | CATGGACTCT | PlateH_F10 | i5MUDI-750 | AGGCATTCTA | TAGAATGCCT
PlateH_G10 | i7MUDI-751 | CAACTCTTAG | PlateH_G10 | i5MUDI-751 | CAACGAATCG | CGATTCGTTG
PlateH_H10 | i7MUDI-752 | GCGTGATTGG | PlateH_H10 | i5MUDI-752 | ACGTTAGCAA | TTGCTAACGT
PlateH_A11 | i7MUDI-753 | CGTAAGACAT | PlateH_A11 | i5MUDI-753 | AATGTCCTGG | CCAGGACATT
PlateH_B11 | i7MUDI-754 | ACACACCTCG | PlateH_B11 | i5MUDI-754 | CGTAAGCTAT | ATAGCTTACG
PlateH_C11 | i7MUDI-755 | TCACGTAGTA | PlateH_C11 | i5MUDI-755 | ACAGCCTGAG | CTCAGGCTGT
PlateH_D11 | i7MUDI-756 | TGAGGCACAT | PlateH_D11 | i5MUDI-756 | CAACACGCGA | TCGCGTGTTG
PlateH_E11 | i7MUDI-757 | GCTAGGATAG | PlateH_E11 | i5MUDI-757 | CGCTGTGGAT | ATCCACAGCG
PlateH_F11 | i7MUDI-758 | ATTCTCGTTG | PlateH_F11 | i5MUDI-758 | ACCATGCTGT | ACAGCATGGT
PlateH_G11 | i7MUDI-759 | CGTACATCCA | PlateH_G11 | i5MUDI-759 | CGAGACGTAG | CTACGTCTCG
PlateH_H11 | i7MUDI-760 | ACTGTCTCAT | PlateH_H11 | i5MUDI-760 | ACCTCCTTAA | TTAAGGAGGT
PlateH_A12 | i7MUDI-761 | AACGGAGTCG | PlateH_A12 | i5MUDI-761 | CGATAACGCG | CGCGTTATCG
PlateH_B12 | i7MUDI-762 | ACGCTGGTAG | PlateH_B12 | i5MUDI-762 | ACAGGTGTTG | CAACACCTGT
PlateH_C12 | i7MUDI-763 | CGTGTCTCTA | PlateH_C12 | i5MUDI-763 | AGTCGCACAG | CTGTGCGACT
PlateH_D12 | i7MUDI-764 | AGACGATGCT | PlateH_D12 | i5MUDI-764 | AGCATACGGA | TCCGTATGCT
PlateH_E12 | i7MUDI-765 | CCTATCACGG | PlateH_E12 | i5MUDI-765 | ACATCGATGG | CCATCGATGT
PlateH_F12 | i7MUDI-766 | ACCATAGGTA | PlateH_F12 | i5MUDI-766 | AACGGCTCAC | GTGAGCCGTT
PlateH_G12 | i7MUDI-767 | CGGTAATGTT | PlateH_G12 | i5MUDI-767 | AACCACAGTT | AACTGTGGTT
PlateH_H12 | i7MUDI-768 | CGTCGTCAAG | PlateH_H12 | i5MUDI-768 | CCAACGACAG | CTGTCGTTGG